Tulip Medical

Quality System Procedure

QSP 7.5-2

Identification and Traceability

Approvals

VP Regulatory Affairs	Date

Regulatory Affairs Coordinator	Date

Document Control Administrator	Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 05/06/2013 | Initial release. |
| 1 | 10/23/2017 | References to NetSuite substituted with references to SOS Inventory software. Updated for accuracy. |
| 2 | 08/29/2018 | Updated for requirements for unique identifiers and identification through the production process. |
| 3 | 10/01/2019 | Section 2.0 Scope – update paragraph Section 6.0 Responsibilities – updated the “Preservation” paragraph Section 7.5 Store Finished Product – added containers and location Section 7.6 Send Product to Customer – 7.6.2 updated paragraph |
| 4 | 02/23/2021 | Added section 7.3.2 Labeling – added reference QSP 7.5-3 |
| 5 | 5/23/2025 | Updated labeling section to include specific reference to MDR 2017/745 Annex VI Part C for Unique Device Identification (UDI) requirements to support MDR Stage I audit readiness. No references to MDD were found in this document. |

- PURPOSE

The purpose of this procedure is to establish product identification and traceability controls for use by Tulip Medical Products.

- SCOPE

This procedure applies to all devices, whether manufactured by Tulip or by Contracted Suppliers.

- REFERENCE DOCUMENTS

| Form 6.2-1-2 | Training Record |
| --- | --- |
| Form 7.4-2-0-1 | SOS Inventory Receiving IIR Log |
| Form 7.4-2-1 | Receiving Inspection Form |
| Form 8.3-1-1 | Nonconformance Report |
| QSP 6.2-1 | Training and Qualification |
| QSP 7.4-2 | Product Inspection |
| QSP 7.5-1 | Device History Record |
| QSP 8.3-1 | Control of Nonconforming Product |
| QSP 7.5.3 | Product Labeling |

- DEFINITIONS

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (QSPs).

- ABBREVIATIONS

DHR – Device History Record

MR – Management Representative

PM – Production Manager

QC – Quality Control

- RESPONSIBILITIES

The Management Representative (MR) is responsible for maintaining this procedure in a current state, and for ensuring that Tulip Medical employees whose jobs encompass these procedures follow them accordingly.

The Production Manager (PM) is responsible for the following:

            - Identification: PM identifies product status throughout production, storage, installation and servicing to ensure that only product that has passed the required inspections and tests, or released under an authorized concession, is dispatched, used or installed.

            - Preservation: PM overseas the preservation of product conformity to requirements during processing, storage, handling and distribution. Preservation is applied to the constituent parts of a medical device. They are responsible for protecting product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling and distribution by:

  - Designing and constructing suitable packaging and shipping containers;

  - Documenting requirements for special conditions needed if packaging alone cannot provide preservation

If there is a new regulation and documenting it is required, Quality Control shall document a system to assign unique device identification to the medical device.

- PROCEDURE

  - Receive Material at Tulip Medical Products

    - Receive material according to QSP 7.4-2 Product Inspection.

    - Tulip Medical Products requires its suppliers to provide sufficient traceability information. These include, but are not limited to

      - Packing lists

      - Purchase order references

      - Lot numbers

      - Serial numbers

    - Traceability information is recorded on Form 7.4-2-1 Receiving Inspection Form and referenced in Form 7.4-2-0-1 SOS Inventory Receiving IIR Log as part of the Device History Record (DHR). The initial inspection forms are filed in accordance with QSP 7.5-1 Device History Record.

    - The DHR file folder includes the product IFU, product labels, and Form 7.4-2-1 Receiving Inspection Form.

  - Maintain Materials in Storage

    - All components and finished product held in storage have traceability information identifying the product. The location of product is referenced on Form 7.4-2-1 Receiving Inspection Form.

  - Manufacturing / Assembly / Packaging / Labeling Activities

    - Maintain identification and traceability information throughout all manufacturing, assembly and packaging operations. SOS Inventory software maintains traceability records pertaining to suppliers, purchase orders and material component subparts used in the assembly of product or received finished product.

    - Labeling – Ensure all labels comply with QSP 7.5-2 and include any applicable UDI requirements as per Regulation (EU) 2017/745, Annex VI, Part C.

  - Production Complete

    - Retain identification and traceability information for finished lots in the appropriate DHR.

      - See QSP 7.5-1 Device History Record for more information.

  - Store Finished Product

    - Store finished product in clearly labeled containers in marked locations and maintain traceability information.

  - Send Product to Customer

    - All product sent to customers has the associated lot and/or serial numbers identified on the shipping documents.

    - Retain shipping traceability information in case of a recall. Documents with lot and/or serial numbers can be easily traced.

  - Disposition of Product

    - Handle any nonconforming product according to QSP 8.3-1 Control of Nonconforming Product.

    - If applicable, file a copy of Form 8.3-1-1 Nonconformance Report with the DHR for traceability purposes.

- FLOWCHART

Store Finished Product

Disposition Product

Send Product to Customer

Production Complete

Manufacturing / Assembly / Packaging Activities

Maintain Materials in Storage

Receive Material