Tulip Medical

Quality System Procedure

QSP 7.6-1

Control of Measurement Equipment

Approvals

VP Regulatory Affairs	Date

Regulatory Affairs Coordinator	Date

Document Control Administrator	Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | N/A | Initial release. |
| 1 | 11/06/2013 | Step 3 – Added – “Note: Calibration may be done on any day of the month the calibration is due. For example, if the calibration frequency of any instrument is annual and the calibration was done on the first of June, the calibration is due by June 30th of the following year.” |
| 2 | 09/26/2018 | Reformatted and included new signature blocks. |
| 3 | 10/18/2018 | Inclusion of Form 7.3-4-4 Installation Qualification for Equipment Inspection Maintenance and Calibration. Updated Section 8.0 Flowchart. |
| 4 | 07/10/2020 | Added requirements for nonconforming equipment (7.14.4); added requirements for protecting equipment from damage (7.14.2); changed font and font size |
| 5 | 3/14/2023 | Changed Equipment Qualification Record form number to form 6.3-1-3 in reference section and section 7.2 to match actual form number. Added form 7.6-1-1 Calibration Master List. Added form 7.6-1-2 Equipment Calibration Record Added Clarity to Sections 7.4 (a), 7.10.2 (b), 7.5.1 (c), 7.14.2.1 (d), 7.14.2 (e), to show Intent of ISO 13485:2016 Clause 7.6 a – e – IA finding ID#6 and CAPA 22-020-IA |
| 6 | 5/23/2025 | Added MDR alignment reference to EU MDR 2017/745, Annex I, Section 14.2 to ensure conformity with GSPR requirements related to measurement equipment. |

- PURPOSE

To describe the calibration system and how calibration records are documented at Tulip Medical Products.

- SCOPE

This Quality System Procedure is to be followed by all Tulip Medical Products personnel. Comprehension of, and satisfactory expertise in, the material covered by this Quality System Procedure by each employee will be indicated through the completion of all applicable and appropriate training requirements, as is indicated by the completion of Form 6.2-1-2 Training Record in accordance with QSP 6.2-1 Training and Qualification.

- REFERENCE DOCUMENTS

EU MDR 2017/745, Annex I – Section 14.2: Devices shall be designed and manufactured in such a way as to ensure the characteristics and performance requirements of the measuring and control equipment used for monitoring and testing are suitable for their intended purpose and are regularly verified.

| Form 4.2-1-1 | Change Request Log |
| --- | --- |
| Form 4.2-1-2 | Change Request |
| Form 6.2-1-2 | Training Record |
| Form 7.1-1-1 Form 6.3-1-3 | Project Control Form Installation Qualification Record |
| Form 7.6-1-1 | Calibration Master List |
| Form 7.6-1-2 | Equipment Calibration Record |
| QSP 4.2-2 | Control of Quality Records |
| QSP 5.6-1 | Management Review |
| QSP 6.2-1 | Training and Qualification |
| QSP 7.1-1 | Risk Management |
| QSP 7.3-2 | Device Master Record |
| QSP 8.2-1 QSP 8.3-1 | Quality Audits Control Of Nonconforming Products |
| QSP 9.1-3 QSP 9.5-1 RSK 04 RSK 07 RSK 08 RSK 09 | Development of Technical File Advisory Notices, Recalls & Vigilance Process FMEA for Disposables Process FMEA for Reusables Design FMEA for Reusables Design FMEA for Disposables |

- DEFINITIONS

Calibration – A set of operations that compare values taken from a piece of inspection, measuring and/or test equipment, or a gauge to a known standard under specified conditions.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (QSPs).

Verifications – Confirmation by examination and provision of objective evidence that specified requirements have been fulfilled.

- ABBREVIATIONS

ID – Identification Number

MR – Management Representative

NIST – National Institute of Standards and Technology (US)

NCR – Non-Conformance Report

- RESPONSIBILITIES The Management Representative is responsible for maintaining this procedure in its current state.

Production Manager is responsible for ensuring that the external resources used to calibrate equipment have the necessary skills, equipment and training.

Production and Production Manager are responsible for ensuring that the tools and equipment that are used for measuring product have been verified as acceptable for use.

- PROCEDURE

The organization shall document procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements

  - Identify Measuring and Test Equipment

  - For each new piece of equipment and/or calibrated instrument, the MR completes Form 6.3-1-3 and the completed Form becomes a Record that is maintained in the Equipment Maintenance folder in the Company’s QMS.

    - All equipment used to inspect and test components, or products, are identified individually with an Identification Number (ID) and recorded on Form 7.6-1-1 Calibration Master List.

    - Any measuring or test equipment that does NOT require calibration (meaning it is not used to control components or product), must be labeled as ‘Not Calibrated’ and/or ‘For Reference Only.’

  - Assign a Unique ID Number

    - Each tool and/or equipment that is under calibration control and does not have a model and serial number attached to it must have a unique and permanent identifier assigned to it.

    - A list of all monitoring and measuring equipment, including their unique ID, is maintained by Operations. Refer to the Calibration Status Record and Log.

    - Calibration records are maintained for each piece of monitoring and test equipment.

    - Records shall remain legible, readily identifiable and retrievable. Changes to a record shall remain identifiable

  - Determine Frequency and Source of Calibration Determine the frequency of calibration for each piece of equipment requiring calibration. Can be calibrated or verified, or both, at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards: when no such standards exist, the basis used for calibration or verification shall be recorded

Note: This can be based on recommendations from the equipment supplier.

Note: All new test equipment should come with a certificate of calibration. This certificate is to be placed in the calibration records for that piece of equipment and the date the equipment is placed into service can be recorded as the calibration date on the Calibration master List.

    - Document the calibration date and the next calibration due date on the Calibration Master List.

Note: Calibration may be done on any day of the month the calibration is due.

  - Equipment Identification

    - Each piece of equipment requiring calibration shall have a calibration label indicating the date of calibration and the next calibration due date.

  - Has the Due Date Been Reached?

    - The Production Manager will monitor the tools/equipment calibration due dates.

    - When the due date has been reached, the item must be obtained and calibration performed, or the equipment replaced.

Note: All calibrations are to be performed by the calibration due date. If the item has not been calibrated by this date, then the equipment is to be tagged with “DO NOT USE – OUT OF CALIBRATION” label or the equipment removed from Tulip Medical Products.

  - Internal Calibration

    - Determine if the measuring tool can be calibrated internally and follow the appropriate step.

Note: At, or before an instrument’s calibration due date, Tulip Medical Products may decide to purchase new equipment rather than calibrate the equipment. In such cases, records of calibration of the new equipment will be maintained, as necessary, and any equipment with an expired calibration due date will be discarded or properly identifies as out of calibration.

  - Calibrate Measuring Tool

    - Calibrate measuring tools in accordance with Calibration Work Instructions, if calibrated in house.

    - Record ‘Master Gauge’ used for calibration. This provides traceability to external calibration and standard complied with.

    - If calibration is successful, go to 7.10 of this procedure.

    - If the calibration revealed that the equipment was out of specified limits, the Production Manager is to be immediately notified so that the possible effect on quality on any produced parts may be evaluated.

    - If there is an impact on material/product, the Production Manager must take appropriate steps in disposition of the non-conforming pieces, including completion of an NCR per QSP 8.3-1 Control of Nonconforming Product.

  - Use Suitable External Calibration Source

    - Determine external source for calibration and ensure source is on Approved Suppliers List.

Note: External calibration means the source of the calibration effort.

    - Monitoring measuring devices may be purchased already calibrated by the manufacturer, along with the device calibration records. By the calibration due date, these instruments are either recalibrated or replaced with newly calibrated instruments.

  - Obtain Calibration Certificate as Record

    - The calibration record must show traceability to national (NIST) or international standards. Before NIST, it could be traceable to another ISO Standard.

    - If the calibration was successful, or must be adjusted or re-adjusted, such adjustments or re-adjustments shall be recorded.

    - If the calibration revealed that the equipment was out of specified limits, the Production Manager is to be immediately notified so that the possible effect on quality of any produced parts may be evaluated.

    - If there is an impact on material/product, the Production Manager must take appropriate steps in disposition of the non-conforming pieces, including completion of a NCR per QSP 8.3-1 Control of Nonconforming Product.

  - Records Maintained

    - Following calibration, the results are to be recorded on the Calibration Master List. Calibration records/certificates are to be filed in the appropriate equipment file. Records shall be protected from damage and deterioration during handling, maintenance and storage The controlled copy of the Calibration Master List is maintained in the Company’s Calibration Folder per QSP 4.2-2 Control of Quality Records. The calibration sets in the Product Receiving / Inspection Room, in the same drawer. The Production Manager is responsible for maintaining this record.

  - New Due Date Determined

    - The next due date is based on a technical assessment of consequence of error, results of most current calibration, the historical records of previous calibration events, and/or the manufacturer’s recommendations.

    - If readings were not within acceptable value, the Production Manager will take actions as described in 7.7 of this procedure. The interval may be decreased based on the above points should any negative indication be found.

    - Should there be a change in calibration interval, this is to be noted on the Calibration Status Record and Log for the item with the effective date noted.

  - New Calibration Label Placed on Item

    - Following successful calibration, the calibration label is to be affixed to the item to signify that the equipment has been verified and should remain within specified tolerances until the next due date.

    - This calibration label will be either the label supplied by the external calibration source, or a Tulip Medical products-supplied calibration label.

    - The calibration label is affixed to the item to indicate:

- Date of calibration

- Due date of the next calibration

- Initials of the person that calibrated the equipment

  - Continue Use

    - The equipment can be placed back into use and continued to be used, until the next calibration due date.

    - All equipment shall be protected from damage during handling, maintenance, and storage.

      - Handle all equipment with care during use. Use equipment according to its user manual if applicable.

      - Equipment shall be handled with care during maintenance and be safeguarded from adjustments that would invalidate the measurement result. Maintain equipment according to its user manual if applicable.

      - Store all equipment in a designated area as to minimize damage/exposure. Small equipment shall be stored in a protective barrier or case. Large equipment shall be stored away from high traffic zones, and away form other objects that may fall onto the equipment.

    - Should at any time the results be questionable, then an unscheduled calibration / verification may be conducted.

    - If equipment is found not to conform to requirements, the validity of previous measurements shall be assessed and recorded.

      - Determine which parts (and corresponding lots) were measured using the equipment in question

      - Pull samples from the lot(s) in question and take new measurements with calibrated equipment.

      - If new measurements match the original measurements (see part acceptance criteria for tolerances), document the results and add to the DHR. No further actions are required.

      - If new measurements do not match the original measurements (see part acceptance criteria for tolerances), document the results and add to the DHR.

        - Determine the severity of any hazards that may occur due to the measurement being out of tolerance (see RSK 04, RSK 07, RSK 08, and/or RSK 09).

        - Use QSP 9.5-1 to determine if any advisory notices, recalls, and/or vigilance plans are required.

- PROCESS FLOWCHART