Tulip Medical

Quality System Procedure

QSP 9.1-2

Conducting Recalls – US FDA

Approvals

11/12/2018

Director of Operations 						Date

11/12/2018

Management Representative 						Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 05/08/2013 | Initial release. |
| 1 | 09/11/2018 | Reformatted. Approval block changed to require approval from Regulatory Affairs. |
| 2 | 11/06/2018 | Revised sections 3.0 and 7.5.1 to read “Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan.” |

- PURPOSE

To describe the procedure used in case of a product recall or product retrieval of any product manufactured and/or distributed by Tulip Medical Products. This includes details for definition, initiation, strategy, and implementation and follow up of a product recall.

- SCOPE

These procedures apply to all situations which could conceivably result in a recall, product retrieval action, product correction, market withdrawal, and where the FDA requires contacting in line with 21 CFR Part 806.

Tulip Medical Products may voluntarily initiate a recall, or the FDA may require it.

- REFERENCE DOCUMENTS

| Food and Drug Administration 21 CFR Part 7, Subpart C – Recalls (Including Product Corrections) – Guidance on Policy, Procedures, and Industry Responsibilities |  |
| --- | --- |
| Food and Drug Administration 21 CFR 806 Medical Devices; Reports of Corrections and Withdrawals |  |
| Form 6.2-1-2 | Training Record |
| Form 8.2-2-2-2 | Complaint Investigation |
| Form 9.1-2-1 | Health Hazard Evaluation Form |
| Form 9.1-2-2 | Product Recall Worksheet |
| Form 9.1-5-1 | Advisory Notices, Recalls and Vigilance Plan |
| QSP 6.2-1 | Training and Qualification |
| QSP 8.2-3 | Customer Feedback / Complaints |
| QSP 9.1-5 | Advisory Notices, Recalls, and Vigilance |
| QSP 9.1-6 | Medical Device Reporting |

- DEFINITIONS

Complaint – Any written, electronic, or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety, effectiveness, or performance of a device after it is released for distribution.

Correction – A repair, modification, adjustment, relabeling, destruction, or inspection (including patient monitoring) of a product without its physical removal to some other location.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (QSPs).

Market Withdrawal – A firm’s removal or correction of a distributed product that involves a minor violation that would not be subject to legal action by FDA, or which involves no violation, e.g. normal stock rotation practices or technology upgrades.

Medical Device Report (MDR) – A US FDA regulation (21 CFR 803) that requires manufacturers to report to the FDA incidents where death or serious injury, or their potential, may have or could have occurred. Refer to QSP 9.1-6 Medical Device Reporting for additional information.

Recall – Recalls are conducted when there is a risk of death or serious deterioration to the state of health caused, or contributed to, by a device. Recall means a firm’s removal or correction of a marketed product that the FDA considered to be in violation of the laws it administers and against which the FDA would initiate a legal action.

- E.g. seizure, the return of a medical device to the manufacturer, its modification y the manufacturer at the site of installation, its exchange, or its destruction.

Recall Classification – The numerical designation (i.e. I, II, or III) assigned by the FDA to a particular product recall to indicate the relative degree of health hazard presented by the product being recalled. Classifications are listed below.

- Class I – A situation in which there is a reasonable probability that the use of, or exposure to, a violative product will cause serious adverse health consequences or death.

- Class II – A situation in which use of, or exposure to, a violative product may cause temporary or medically reversible adverse health consequences or where the probability of serious adverse health consequences is remote.

- Class III – A situation in which use of, or exposure to, a violative product is not likely to cause adverse health consequences.

Recall Strategy – A planned specific course of action to be taken in conducting a specific recall, which addresses the depth of the recall, need for public warnings and extent of effectiveness checks for the recall.

Recall Team – At a minimum includes the Management Representative (MR) and at least one other member of the executive-level staff.

Remedial Action – Any action other than routine maintenance or servicing of a device where such action is necessary to prevent recurrence of a reportable event.

Stock Recovery – A firm’s removal or correction of a product that has not ben marketed or that has not left the direct control of the firm, e.g. the product is located on premises owned by, or under control of, the firm. And no portion of the lot has been released for sale or use.

- ABBREVIATIONS

CFR – Code of Federal Regulations (US)

EVP – Executive Vice President

FDA – Food and Drug Administration

MDR – Medical Device Report

MR – Management Representative

RA – Regulatory Affairs

- RESPONSIBILITIES

The Management Representative is responsible for recommending the Recall Action Plan and for assuring implementation of the plan once it is approved.

The Executive Vice President (EVP) is responsible for reviewing the Recall Action Plan and approving it before the plan is initiated.

Regulatory Affairs (RA) communicates with Regulatory Authorities to review any proposed market withdrawal or stock recover action.

- PROCEDURE

Note: Refer to CFR 21 Part 7, Subpart C and Part 806 for guidance on device recalls. If the determination is made to recall product(s), consult the current Regulatory Consultant and/or legal counsel for assistance.

    - General Notes

    - Those adverse incidents involving particular issues of significant public health concern should be reported regardless of exemption criteria.

    - Form 9.1-2-1 Health Hazard Evaluation Form and Form 9.1-2-2 Product Recall Worksheet may be used, but are not required.

  - Using CFR 21 Part 806, create the list of tasks and data to be gathered. For example:

- Identify/document the product involved – Reference Number; Lot Number(s); Serial Number(s); Quantities; Date(s) released to Finished Goods.

- Document the problem, including an assessment of the health hazard.

- Put together a Recall Team.

- Identify the customers that have received the product in question.

- Identify the means of customer communication, including response to the FDA.

- Identify how the product is to be returned, as well as how it is to be accounted for when returned.

- Identify the disposition of returned goods.

  - Situation Identification

    - Problem Definition: State exactly what the problem is.

Note: A potential problem would be a situation in which the product falls into a violative product category.

    - The extent of the problem likely to occur again; how many products have been distributed and to what level; and where the product is now.

    - Evaluate the problem for impact on product performance and functionality.

    - The health hazard and likelihood of injuries (risk) associated with the problem.

      - Evaluate any adverse affects that the problem may introduce to a patient, taking into consideration serious harm, complications, or death resulting from the problem.

    - The necessity for any immediate internal action, such as putting a hold on distribution of the product, sub-assemblies in process and/or raw material components involved until a final determination is made.

  - Determine if the FDA is to be notified.

    - If applicable, the MR notifies FDA in accordance with CFR 21 Part 806.

Note: Contact the regulatory consultant and/or legal counsel for assistance.

    - If NO notification is to be made, document this decision and maintain this documentation with any applicable Complaint and/or Investigation paperwork.

    - The decision for a market withdrawal, or no recall, rationale must be documented to show if a minor violation has occurred, but no FDA notification is required.

  - Recall Action Plan

    - If the recommendation is to recall the product based on the above information, or correct the product in the field, the MR shall prepare the written proposal for action on Form 9.1-5-1Advisiory Notices, Recalls and Vigilance Plan, which is to be provided to the EVP and RA for review and approval.

    - The Recall Action Plan must include the following information:

- Identify the problem, describing how knowledge of the problem occurred and the subsequent follow-up activities.

  - It should include actual numbers and results of any analysis and/or testing performed.

- Scope of the action, recall, or product correction.

- An evaluation and conclusion stating the possible patient hazards that may be associated or arise from the product.

  - This is performed and documented on Form 9.1-2-1 Health Hazard Evaluation Form.

- Specify the depth of the recall (the level in the distribution chain to which the recall is to extend).

- Determination of a need for public warning. Special provisions apply when this is felt to be necessary (reference 21 CFR § 7.42 (b) (2).

- Determination of the appropriate effectiveness check level.

- List actual Lot Number(s) and corresponding amount(s).

- Determination of the type and content of the recall communication to be used.

- Determination of the recall implementation method, such as fax, first class certified mail, or telephone call.

- Determination of any in-field product correction. This includes description of the method of correction, who will be responsible for performing the field correction, how the correction will be communicated to the customer, and how the correction(s) will be verified.

- Possible class of recall or product correction.

- Proposed disposition of the recalled or corrected product.

    - The EVP shall review the proposed Recall Action Plan with RA and advise the MR of the final decision.

    - The MR assures implementation of the approved Recall Action Plan, including coordination and implementation of the plan with domestic and international distributors, as appropriate.

  - Prepare customer communication and notify the affected customers using the standard format to include all the necessary information. Include distributors that are not directly associated with the recalled product to make them aware of the situation.

  - The MR notifies the appropriate FDA district office within fifteen (15) working days of when the recall or product correction has been implemented unless specifically stated that no recall notification is to be conducted.

    - Provide the FDA with the necessary information of the recall or product correction, but do not include extraneous information not related directly to the recall.

  - All returned product is to be identified as nonconforming and segregated. Product disposition will be in accordance with the determination made above.

  - Keep a written record of any product returned (source and quantity) or corrected in the field (source, method and quantity) and include this data in the recall file.

  - The FDA may require witness of the actual product disposal. Notify the appropriate FDA office when the recall has been completed and when a reasonable amount of the product are expected to be returned.

  - The MR is responsible for monitoring Recall action, including customer response and product returns, until Recall is determined to be closed.

  - Provide periodic reports to the FDA and/or Recall Team to keep them informed of the status of the recall/withdrawal.

Note: These reports should be completed monthly, or as stated by the Recall Action Plan.

  - When all reasonable efforts have been made to remove or correct the product in accordance with the action plan and commensurate with the degree of the hazard, the recall shall be completed.

Note: The Corrective Action must be reviewed by the EVP and verified against the intended action listed in the Recall Action Plan.

  - The following information must be retained in the recall file:

- A copy of the Complaint Record and/or Nonconformance Report.

- Documentation of follow-up and attempts to follow-up, and obtain additional information about the event.

- An explanation of why certain information cannot be obtained.

- Copies of any relevant test reports and investigation reports.

- Documentation related to decisions used to determine the reportability of the event.

- Documentation of the final assessment of the event and any corrective or remedial action taken.

- Copies of reports and other information submitted to the applicable authorities.

- Copies of information submitted to distributors or customers.

- Copies of any other information required according the regulations.

  - The close-out shall be in the form of a Final Status Report to the EVP. The MR shall also review the company’s file of the completed action to assure that all appropriate documentation is in place and to assure that no extraneous information is included.

  - The Final Status Report shall be sent to the FDA notifying them that the recall has been completed. A copy of the Final Status Report submitted to the FDA shall be maintained in the recall file.

  - If the FDA verbally requests additional information, the MR should ask the FDA to follow up such requests in writing.

- PROCESS FLOWCHART

  - N/A