Approval

V.P Regulatory Affairs 						             Date

Regulatory Affairs Coordinator					             Date

Document Control Administrator					              Date

| Rev | CR # | Effective / Approval Date | Description of Change | Originator |
| --- | --- | --- | --- | --- |
| 0 | 21-039 | 3/27/2021 | Release documentation - Implementation of document for EU Medical Device Regulations (MDR) 2017/745 for Class I Self-Certified Devices | Marc Pilkington |
| 1 | 25-038E | 03/28/2025 | Corrected various spelling errors and typos. | Marc Pilkington |
| 2 | 25-061 | 07/28/2025 | Clarifications and updates per MDR/ISO/MEDDEV including state of the art, PRISMA, PMCF justification, CER triggers, and literature review practices. Added Appendix E - Sample CER Authoring Checklist | Marc Pilkington |

- PURPOSE -- SCOPE

The purpose of this procedure is to describe the process to be used by Tulip Medical Products to provide clinical data in support of the clinical safety and performance of its device(s) and to meet the clinical evaluation requirements of the Medical Devices Regulation (MDR) (EU) 2017/745 (Article 61 and Annex XIV, Part A).

The procedure provides an outline of the clinical evaluation processes, including the collection and analysis of data, and describes how to compile a Clinical Evaluation Report (CER), which is part of the technical documentation used by Tulip Medical Products to obtain CE marking in accordance with the MDR.

The clinical evaluation process is described in Annex XIV, Part A of the MDR and may be summarized as:

- Establish/update the Clinical Evaluation Plan (CEP);

- Identify available clinical data relevant to the device and its intended purpose and any gaps in clinical evidence through a systematic literature review;

- Appraise all relevant clinical data by evaluating its suitability for establishing the safety and performance of the device;

- Generate any new or additional clinical data needed to address outstanding issues, through properly designed clinical investigations in accordance with the clinical development plan; and,

- Analyze relevant clinical data to establish conclusions regarding the safety and clinical performance of the device, including its clinical benefits.

Results of the clinical evaluation and its supporting clinical evidence are documented in a CER.  This is an iterative process performed throughout the lifecycle of the device.

- Reference DOCUMENTS

  - EN ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes

  - EN ISO 14155:2011 Clinical Investigation of medical devices for human subjects: Good clinical practice

  - ISO 14971:2019 Medical devices – Application of risk management to medical devices

  - Medical Devices Regulation (MDR) EU 2017/745

  - European Commission MEDDEV 2.7/1 Rev 4 Guidelines on Medical Devices: Clinical Evaluation: A Guide for Manufacturers and Notified Bodies under Directives 93/42/EEC and 90/385/EEC

NOTE: MEDDEV 2.7/1 Rev. 4 is intended for reference; conformity is not stated or implied.  Clinical requirements per the MDR apply to devices subject to the MDR.  A guidance or similar document referencing only the MDD may still be considered state of the art, with the caveat that the wording of the MDR takes precedence when there may be overlapping or conflicting requirements.

  - European Commission MEDDEV 2.12/2 Rev 2 Guidelines on Medical Devices: Post Market Clinical Follow-up Studies, A Guide for Manufacturers and Notified Bodies

  - GHTF SG5/N2R8:2007 Clinical Evaluation

  - IMDRF MDCE WG/N56FINAL:2019 (formerly GHTF/SG5/N2R8:2007) Clinical Evaluation

  - QSP 4.2-1 Document Control

  - QSP 7.3-1 Design and Risk Management

  - QSP 8.2-1 Customer Complaints

  - QSP 9.1-3 Post-Market Surveillance (PMS) and Post-Market Clinical Follow-Up (PMCF)

  - QSP 9.1-2 Technical Documentation - EU MDR

  - Appendix A: Literature Searching Procedure

  - Appendix B: Example Format for Literature Search Report

  - Appendix C: Example Format for CER

  - Appendix D: Flowchart Assessing Citations (Suitability for Inclusion in CER)

- DEFINITIONS

  - Clinical Benefit (MDR, Art. 2(53)): the positive impact of a device on the health of an individual, expressed in terms of a meaningful, measurable, patient-relevant clinical outcome(s), including outcome(s) related to diagnosis, or a positive impact on patient management or public health

  - Clinical Data (MDR, Art. 2(48)): information concerning safety or performance that is generated from the use of a device and is sourced from the following:

- Clinical investigation(s) of the device concerned,

- Clinical investigation(s) or other studies reported in scientific literature, of a device for which equivalence to the device in question can be demonstrated,

- Reports published in peer reviewed scientific literature on other clinical experience of either the device in question or a device for which equivalence to the device in question can be demonstrated,

- Clinically relevant information coming from post-market surveillance, in particular the post-market clinical follow-up.

  - Clinical Evaluation (MDR, Art. 2(44)): a systematic and planned process to continuously generate, collect, analyse and assess the clinical data pertaining to a device in order to verify the safety and performance, including clinical benefits, of the device when used as intended by the manufacturer.

  - Clinical Evidence (MDR, Art. 2(51)): clinical data and clinical evaluation results pertaining to a device of a sufficient amount and quality to allow a qualified assessment of whether the device is safe and achieves the intended clinical benefit(s), when used as intended by the manufacturer.

  - Clinical Investigation (MDR, Art. 2(45)): any systematic investigation involving one or more human subjects, undertaken to assess the safety or performance of a device.

  - Clinical Performance (MDR,  Art. 2(52)): the ability of a device, resulting from any direct or indirect medical effects which stem from its technical or functional characteristics, including diagnostic characteristics, to achieve its intended purpose as claimed by the manufacturer, thereby leading to a clinical benefit for patients, when used as intended by the manufacturer.

  - Device Registry (MEDDEV 2.12/2): an organised system that uses observational study methods to collect defined clinical data under normal conditions of use relating to one or more devices to evaluate specified outcomes for a population defined by a particular disease, condition, or exposure and that serves predetermined scientific, clinical or policy purpose(s).

  - Eudamed: the European database on medical devices that integrates different electronic systems to collate and process information regarding devices on the market and the relevant economic operators, certain aspects of conformity assessment, notified bodies, certificates, clinical investigations, vigilance and market surveillance. {Currently delayed until May 2022}

  - Harmonized Standard (MDR, Art. 2(70)): a European standard as defined in point (1)(c) of Article 2 of Regulation (EU) No 1025/2012 on European standardisation, i.e. a European standard adopted on the basis of a request made by the Commission for the application of Union harmonization legislation.

Note: There are currently no Harmonized Standards for the MDR; in this absence, current Harmonized Standards under the MDD or more recent international versions can likely be used until standards become harmonized under the MDR. Refer to QSP 4.2-1 for monitoring of MDR-related guidance documents and standards.

  - Performance (MDR, Art. 2(22)): the ability of a device to achieve its intended purpose as stated by the manufacturer.

  - Post-Market Clinical Follow-up (PMCF) (Derived from MDR, Annex XIV Part B): PMCF shall be understood to be a continuous process that updates the clinical evaluation and shall be addressed in the manufacturer's post-market surveillance (PMS) plan. When conducting PMCF, the manufacturer shall proactively collect and evaluate clinical data with the aim of confirming the safety and performance throughout the expected lifetime of the device, of ensuring the continued acceptability of identified risks and of detecting emerging risks on the basis of factual evidence.

  - PMCF Plan (Derived from MDR,  Annex XIV Part B): methods and procedures specified by the manufacturer to proactively collect and evaluate clinical data with the objectives to confirm safety and performance of the device throughout its expected lifetime; identify previously unknown side-effects and monitor known side-effects and contraindications; identify and analyze emergent risks as based on factual evidence; ensure continued acceptability of the benefit-risk ratio (reference GSPR 1 and 9); and identify possible systematic misuse or off-label use, with a view to verifying the intended purpose as correct.

  - PMCF Study (Derived from MDR, Art. 74(1)): a clinical study carried out following the  CE marking of a device and conducted to further assess the device within the scope of its intended use.

- ABBREVIATIONS

  - CEP – Clinical Evaluation Plan

  - CER – Clinical Evaluation Report

  - CS – Common Specifications

  - DAEN – Database of Adverse Event Notifications (AUS)

  - EU – European Union

  - FCSA – Field Safety Corrective Action

  - FDA – Food and Drug Administration (US)

  - GHTF – Global Harmonization Task Force

  - GSPR – General Safety and Performance Requirements

  - IMDRF – International Medical Device Regulators Forum

  - MAUDE – Manufacturer and User Facility Device Experience (US)

  - MDR – Medical Devices Regulation (EU)

  - NICE – National Institute for Health and Care Excellence (UK)

  - PMCF – Post-Market Clinical Follow-up

  - PMS – Post-Market Surveillance

  - PRRC – Person Responsible for Regulatory Compliance

  - TD – Technical Documentation

  - TGA – Therapeutic Goods Administration (AUS)

  - WHO – World Health Organization

- RESPONSIBILITIES

  - The PRRC and Regulatory Team are responsible for maintaining this procedure in a current state and are responsible for ensuring that all CEPs and CERs remain current.

  - Tulip Medical Products is responsible for specifying and justifying the degree of clinical evidence necessary to demonstrate conformity with the relevant general safety and performance requirements (GSPR).  The degree of clinical evidence is to be proportionate and appropriate to the device considering its intended purpose, classification, risks and claims. The clinical evaluation shall be thorough and objective, accounting for favorable and unfavorable data.

  - The CER should be prepared by a suitably qualified individual or team knowledgeable in research methodology, information management, regulatory requirements and medical writing.

  - If equivalence to other device(s) is claimed for clinical evaluation purposes, Tulip Medical Products is responsible for demonstrating sufficient levels of access to the data for the other device(s) necessary to justify equivalence.

  - The Regulatory Team and Production are to ensure that all CERs associated with the current Tulip Medical Products’ medical devices marketed and sold in Europe are reviewed on a periodic basis.

- PROCEDURE

  - Background of Requirements

    - Device conformity with the relevant GSPR under normal conditions of intended use, evaluation of undesirable side-effects and benefit-risk ratio acceptability is based on clinical data providing sufficient clinical evidence, including PMS data where relevant.  Clinical evaluation is the process by which clinical data is generated, collected, analyzed and assessed; consequently, clinical evaluation is required for all medical devices.

    - Clinical evaluation is conducted per a defined, methodical process based on:

      - Critical evaluation of available scientific literature related to the safety, performance, design and intended purpose of the device where:

        - The device has been demonstrated to be equivalent to another device with clinical data (refer to Step 6.9), and

        - Compliance with the relevant GSPR is demonstrated.

      - Critical evaluation of all available clinical investigations, considering MDR compliance; and

      - Consideration of alternative treatment options, if available.

    - The clinical evaluation and any literature searches carried out to obtain clinical evidence must be sufficiently documented to ensure that the results can be verified, objectively appraised, and reproduced if necessary (refer to Appendix C in this procedure).

    - If there is insufficient clinical evidence to verify the satisfactory safety and clinical performance of the device, the manufacturer may need to generate additional clinical data (e.g. perform clinical investigations).

    - The manufacturer is obligated to actively update the clinical evaluation throughout the device life cycle with clinical data obtained from their PMCF plan, including conclusions of the PMCF evaluation report and/or PMS plan and report (refer to QSP 9.1-3).

  - Establish/update the CEP

    - Prepare and maintain the CEP to include the following, at minimum:

      - Identify the GSPRs (refer to MDR Annex I) that require support from relevant clinical data (clinical experience/use of the device);

        - For MDR, generally include GSPR 1 and 8.

      - Define the scope of the clinical evaluation based on the GSPRs that need to be addressed from a clinical perspective and risks identified in the device risk management document.

      - Intended purpose of the device;

        - Specify the intended purpose of the device per its labeling.

      - Intended target groups with clear indications and contraindications;

        - Specify the target patient population(s), e.g. medical condition, disease progression, age as well as device indication(s) and contraindication(s) per device labeling.

      - A detailed description of intended clinical benefit to patients with relevant and specific clinical outcome parameters;

      - Specified methods to be used for examining qualitative and quantitative aspects of clinical safety with clear reference to determining residual risks and side-effects;

      - Indicative list and specified parameters to be used to determine, based on the state of the art in medicine, acceptability of the benefit-risk ratio for the various indications and intended purpose(s) of the device;

      - An indication how benefit-risk issues relating to specific components such as use of medicinal products, or non-viable animal or human tissues, are to be addressed; and

      - A clinical development plan indicating progression from exploratory investigations, such as first-in-human studies, feasibility and pilot studies, to confirmatory investigations, such as pivotal clinical investigations, and a PMCF (reference MDR Annex XIV Part B) including an indication of milestones and a description of potential acceptance criteria.

    - The CEP plan will be reviewed and approved by Management Representative at minimum.

  - Identify available clinical data relevant to the device and its intended purpose and any gaps in clinical evidence through a systematic literature review.

    - Review the clinical experience data, clinical investigations and literature reports related to the safety and performance of the subject device and equivalent device(s), if applicable.

    - Perform literature searches to identify the published clinical data for the subject device (e.g., reports of clinical investigations of the subject device performed by third parties, incident reports) and equivalent device(s), if applicable.

Note: Literature search strategies should be reproducible and auditable. Refer to MEDDEV 2.7/1 Rev. 4, Appendix A4 for best practice guidance.

      - Appendix A describes an example of a literature search procedure and includes samples of literature and clinical investigation / registry search sources.

      - Appendix B is an example format for a literature search report.

      - Appendix D contains a flowchart for assessing citations (suitability for inclusion in CER).

      - Appendix E – Sample CER Authoring Checklist

    - Identify pertinent common specifications (CS), harmonized / recognized standards, if extant. Compliance with established standards may help demonstrate that the device technology has established safety and performance characteristics.

Note: MDR Art. 61(10) provides an exception for demonstration of conformity with the GSPR without clinical data based on the manufacturer’s risk management results and specifics of the interaction between the device and human body, intended clinical performance and device claims.  The manufacturer “shall duly substantiate in the technical documentation referred to in Annex II why it considers a demonstration of conformity with general safety and performance requirements that is based on the results of non-clinical testing methods alone, including performance evaluation, bench testing and pre-clinical evaluation, to be adequate”.

    - Compile and Review PMS Data

      - Where the subject device is already marketed, collate and review the manufacturer’s PMS data for the device.

      - Clinical / PMS experience data can be obtained from the following applicable sources:

        - Sales history, complaint history and, if any, incidents;

        - Manufacturer-generated PMS reports, registries or cohort studies;

        - Manufacturer or regulatory authority incident / adverse effect databases;

        - Data from patient compassionate usage programs; and

        - Details from field corrective actions.

      - Identify other similar devices on the market:

        - Length of time commercially available; and

        - Identify whether CE Marked in the EU and commercialized in the United States.

      - Where the clinical evaluation uses clinical data on one or more equivalent devices, search, collect and review available PMS data on the equivalent device(s), including incident / adverse effect reports and FSCAs. Publicly accessible device databases of PMS data include but are not limited to:

        - TGA Database of Adverse Event Notifications (DAEN)

        - US FDA MAUDE, Manufacturer and User Facility Device Experience

        - US FDA Medical Device Recalls

      - Appendix C has a sample table that can be used to identify equivalent and or comparable device(s) for comparison to Tulip Medical Products’ devices.

  - Appraise all relevant clinical data by evaluating in terms of its suitability for establishing the safety and performance of the device.

    - Appraise individual data sets for their suitability and contribution of the results to demonstrate safety and performance of the subject device.

    - Points to consider:

      - Relevance, applicability, quality, clinical significance;

      - Understand merits and limits of clinical data; and

      - Develop appropriate criteria in advance, to be applied consistently for specific circumstances.

  - Generate any new or additional clinical data needed to address outstanding issues, through properly designed clinical investigations in accordance with the clinical development plan.

    - As necessary, conduct a properly designed clinical investigation of the device in accordance with the MDR, clinical development plan per Step 6.2.1.96.2.1.9 and EN ISO 14155 to resolve outstanding issues.

    - The clinical investigation report should include a critical evaluation of data collected during the clinical investigation.

    - Include a summary of the clinical investigation in the CER with references to the clinical investigation report, protocol and other applicable documents and records.

  - Analyze all relevant clinical data to establish conclusions about the safety and clinical performance of the device, including its clinical benefits.

    - Assess whether appraised data sets collectively demonstrate safety and clinical performance of device in relation to intended use.

    - Assess whether the clinical evidence combined with data from non-clinical testing methods and other relevant data address the following:

      - The device performs as intended by manufacturer during normal conditions of use;

      - The device does not compromise the clinical condition or safety of the patient; safety or health of the user, or other persons (where applicable);

      - Risks associated with use of the device are acceptable when weighed against the intended benefits to the patient and are compatible with a high level of protection of health and safety, accounting for the generally acknowledged state of the art; and

      - Known and foreseeable risks and the undesirable side-effects are considered acceptable when weighed against the benefits from performance achieved by the device during normal conditions of use.

  - Consider the need for PMCF. If PMCF is not deemed necessary, provide an evidence-based justification in the CER referencing risk classification, market history, and PMS findings. PMCF is not completely defined in the MDR though appears to be defined as the continuous process that updates the clinical evaluation and includes PMCF investigation, user survey, patient survey, and device registry.  Refer to QSP 9.1-3.

  - Current Knowledge/State of the Art

    - Perform searches to identify the current knowledge / state of the art including applicable standards and guidelines, information on the applicable medical field, condition to be diagnosed or treated with the subject device, benchmark devices, alternative treatments available to the target population.

    - Identify and describe the medical conditions concerned, including different clinical forms, stages and severities if relevant. Discuss the historical context of the device and/or technology as relevant.

    - Describe available alternatives to the subject device (e.g. therapeutic, management, diagnostic) to include risks/benefits/limitations, advantages / disadvantages as relevant and information is available.

    - The current knowledge / state of the art may also be used to support assessments regarding acceptability of benefit-risk ratio and acceptability of side-effects.

  - Device Equivalence

    - Clinical evaluation may be based on clinical data for a device demonstrated as equivalent with respect to technical, biological and clinical characteristics. Characteristics to consider for demonstration of equivalence include:

      - Technical: similar design; similar conditions of use; similar specifications and properties including physicochemical properties, e.g. intensity of energy, tensile strength, viscosity, surface characteristics, wavelength and software algorithms; similar deployment methods, where relevant; similar principles of operation and critical performance requirements;

      - Biological: same materials or substances in contact with the same human tissues or body fluids for a similar kind and duration of contact; similar release characteristics of substances, including degradation products and leachables; and

      - Clinical: same clinical condition or purpose, including similar severity and stage of disease, at the same site in the body, in a similar population, e.g. age, anatomy, physiology; same kind of user; similar relevant critical performance, i.e. expected clinical effect for a specific intended purpose.

    - The clinical evaluation should identify differences between the subject and equivalent device in technical, biological and clinical characteristics.  To be considered equivalent, any differences in the above characteristics shall not result in clinically significant differences in the safety and clinical performance between the devices.

    - Equivalence must be based on scientific justification, which may include comparative test data, labeling or other TD.  To justify its claims of equivalence, Tulip Medical Products shall clearly demonstrate it has “sufficient levels of access to the data relating to devices” for which equivalence is being claimed.  As devices may change over time, Tulip Medical Products is responsible for ensuring any claims of equivalence remain justified over time.

    - The device identified as equivalent should have CE marking or other justification provided.

  - Clinical Evaluation Report

    - The clinical evaluation will be documented as a CER.  See Appendix C for a potential CER format.

    - The CER shall be approved by an Evaluator who should be knowledgeable in the areas listed in Step 5.3 as well as the technology and clinical application of the subject device, diagnosis and management of the condition(s) relevant to the subject device, medical alternatives, treatment standards and technology (e.g. specialist clinical expertise).

    - At minimum a CER Evaluator should have a degree from higher education in the respective field and 5 years of documented professional experience or 10 years of documented professional experience if a degree is not a prerequisite for a given task.

      - Include a copy of the Evaluator’s credentials / Curriculum Vitae as an attachment to the CER.

      - A Declaration of Interest that discloses relevant financial interests is maintained for each current Evaluator.   Declarations should provide the interest, time span included and the degree to which financial interests of family members are included or not (e.g. spouse/partner living in same residence, other persons the evaluator is legally responsible for).  Declarations of Interest are maintained by Financial and Regulatory Departments.

    - The CER is also reviewed and approved by a Tulip executive team member at minimum.

    - The CER will be filed and maintained with the product technical documentation file, e.g. TD section applicable to the ‘Pre-clinical and clinical data’.

  - Periodic Reviews

    - The CEP will be updated throughout the device lifetime when information described in Step 6.2 or otherwise provided by the plan is updated.  The CEP should also be reviewed for potential update when its correlating CER is updated periodically, thereby taking into account the state of the art.

    - The CER will be updated periodically throughout the device lifetime with data obtained from the following sources, as applicable for the product:

      - Clinical data from the PMCF plan implemented by Tulip’s regulatory team for the product and conclusions of the PMCF evaluation report (e.g. additional clinical experience, user feedback, scientific literature review, registries, PMCF studies, clinical data on equivalent or similar devices, new/updated CS, harmonized standards, guidance);

      - Clinical data from PMS plan implemented by Tulip’s regulatory team for the product (e.g. serious incidents including data from PMS reports and FSCAs; non-serious incidents and undesirable side-effects; trend reporting data per MDR Art. 88; specialist or technical literature, databases and/or registries; feedback and complaints; publicly available information regarding similar devices such as regulatory authority adverse effect / vigilance databases; and

      - Resources regarding current knowledge/state of the art (e.g. clinical literature, expert documents produced by professional medical associations such as clinical practice guidelines and consensus statements, CS, harmonized standards).

    - The frequency of CER update is based on whether the device carries significant risks and if it is novel or well-established. Additional CER updates may be triggered by new adverse events, FSCA notices, changes to intended use or claims, or new regulatory guidance. Justification for the frequency of updates is to be documented in the CER.  The CER is updated:

      - Every two to five years if the device is not expected to carry significant risks and is well-established.

    - The CER will also be updated when new PMS information could impact the current evaluation or its conclusions.

- QUALITY RECORDS

  - CEPs and CERs are filed with the product’s TD file.

  - CEPs and CERs are maintained and retained in accordance with QSP 4.2-1.

Appendix A: Literature Searching Procedure

A literature search should be a scientific activity conducted with rigor and objectivity and should allow for verification by third parties. Use of the PRISMA framework (Preferred Reporting Items for Systematic Reviews and Meta-Analyses) is encouraged to ensure transparent and systematic reporting.

- Construct review questions, describe the objective, define the scope (should be consistent with scope of the clinical evaluation). Specify studies relevant to objective considering well established knowledge of the device.

- Develop a search strategy.

Design a protocol to identify, select, and collate relevant publications. Document the Literature Search plan. Identify sources of data, justify the choice, and define the criteria for selecting or excluding data, including published and unpublished data.  Note: Searches for published clinical data shall primarily be conducted in recognized bibliographic databases.

- Develop a database search strategy.

Include a justification for the choice of sources and describe any supplemental strategies (e.g. checking bibliography of articles retrieved, hand searching of literature) used to enhance the sensitivity of the search. Database search details: Search terms (key words, indexing headings) and their relationships (Boolean logic) and the medium used (e.g. online, CD-ROM [include publication date and edition]). Attach a copy of the downloaded, unedited search strategy, and the selection criteria used to choose articles.

- Conduct searches and document results.

Perform literature search per Literature Search plan.  Review results per established selection criteria to identify relevant articles.

- Compile report.

Include full-text copies of relevant published articles for the period covered by the search.

Examples of Literature Sources

Searches for published clinical data shall primarily be conducted in recognized bibliographic databases such as:

    - PubMed www.ncbi.nlm.nih.gov/pubmed

    - Embase www.embase.com

    - NICE www.nice.org.uk

    - The Cochrane Library www.cochranelibrary.com

Clinical trial registers sources may include, but are not limited to:

- US National Library of Medicine ClinicalTrials.gov www.clinicaltrials.gov

- WHO International Clinical Trials Registry Platform www.who.int/ictrp/en/

- CenterWatch www.centerwatch.com

- EU Clinical Trials Register www.clinicaltrialsregister.eu/ctr-search/search

Other sources of literature include:

- Scientific literature databases

- Internet searches/resources for information concerning:

  - Harmonized standards or other applicable standards

  - Field safety corrective actions for equivalent and/or other devices which can be found on manufacturer, European Competent Authority and US Food and Drug Administration sites, and possibly other sites

  - Implant registry reports

  - Documents in systematic review databases (e.g. Cochrane Database of Systematic Reviews, Propsero international prospective register of systematic reviews)

  - Clinical practice guidelines and consensus statements published by professional medical associations

  - Review and meta-analyses prepared by health technology assessment institutes and networks

  - Clinical studies (e.g. WHO International Clinical Trials Registry Platform)

- Non-published data

- References listed in primary sources

Appendix B: Possible Format for Literature Search Report

The sections below provide a potential Literature Search Report format derived from IMDRF MDCE WG/N56FINAL:2019 (formerly GHTF/SG5/N2R8:2007), Appendix B.

- Device name/model

- Scope of the literature search

- Methods

- Dates(s) of search

- Name of person(s) undertaking the literature search

- Period covered by search

- Literature sources used to identify data

- Scientific databases: Bibliographic databases, specialized databases

- Systematic review databases

- Clinical trial registers

- Adverse event report databases

- Reference texts

- Database search details

- Search terms (key words, indexing headings) and relationships (Boolean logic)

- Medium used (e.g. online, CD-ROM (incl. publication date and edition))

- Inclusion/exclusion criteria used to select articles

- Outputs

- Literature citations retrieved from each database search

- Data selection process

Appendix C: Possible Format for CER

IMDRF MDCE WG/N56FINAL:2019 (formerly GHTF/SG5/N2R8:2007), Appendix G and MEDDEV 2.7/1, Appendix A9 provide possible formats for the CER.  The section titles below provide a possible CER format derived from MEDDEV 2.7/1, Appendix A9.  MEDDEV 2.7/1 Appendix A9 contains a potential example of content to be included in each numbered section, except section 4.6.

- Summary

- Scope of clinical evaluation

- Clinical background, current knowledge, state of the art

- Device under evaluation

  - Type of evaluation

  - Similar devices

  - Equivalent device (only for when equivalence is claimed) [See Table 1 for sample]

  - Data generated and held by the manufacturer

  - Data from literature

  - Data from PMS activities

  - Summary and appraisal of data

  - Analysis of the data

    - Requirement on safety (GSPR 1)

    - Requirement on acceptable benefit-risk profile (GSPR 1)

    - Requirement on performance (GSPR 1)

    - Requirement on acceptability of side-effects (GSPR 8)

- Conclusions

- Date of next clinical evaluation

- Dates and signatures

- Qualification of the responsible evaluators

- References

Appendices

Literature Search Protocol and Report

PMCF Evaluation Report

Curriculum Vitae of CER Evaluator(s)

Declaration of Interest of CER Evaluator(s) (unless maintained in other location)

Copies of full-text publications

Table 1 – Sample of Equivalence Determination

| Criteria for Comparison | Subject Device [Manufacturer, Trade Name] | Equivalent Device [Manufacturer, Trade Name] | Rationale for Claiming No Clinically Significant Difference in Device Safety or Performance |
| --- | --- | --- | --- |
| Clinical |  |  |  |
| Clinical condition or purpose / indication |  |  |  |
| Anatomical location in the body |  |  |  |
| Patient population (including age, anatomy, physiology) |  |  |  |
| Similar critical performance in view of expected clinical effect for a specific intended purpose |  |  |  |
| Technical |  |  |  |
| Conditions of use |  |  |  |
| Specifications and properties (e.g. tensile strength, viscosity, surface characteristics, energy intensity) |  |  |  |
| Design |  |  |  |
| Principles of operation and critical performance requirements (where relevant) |  |  |  |
| Biological |  |  |  |
| Materials or substances in contact with tissues or body fluids |  |  |  |
| Duration of contact with tissues or body fluids |  |  |  |
| Release characteristics of substances including degradation products and leachables |  |  |  |

Appendix D: Flowchart Assessing Citations (Suitability for Inclusion in CER)

The below flowchart visually outlines the process used in assessing citations retrieved from queries of online databases for suitability for inclusion in the CER.

Review titles and abstracts of search results against inclusion/exclusion criteria.

Obtain article and evaluate in detail against inclusion/exclusions criteria.

Relevant literature subjected to the appraisal process and inclusion as appropriate in the CER.

Meets criteria?

Meets criteria?

YES

YES

NO

NO

Exclude from report with rationale.

Exclude from report with rationale.

END OF DOCUMENT

State of the Art: The accepted standard of care or available treatment options at a point in time, based on clinical evidence, best practices, and regulatory expectations.

# Appendix E – Sample CER Authoring Checklist

CER Authoring Checklist – Focused Verification (QSP 9.4-1 Template)

Document: __________		Revision: ____

| Checkpoint | Verification | Initials |
| --- | --- | --- |
| PRISMA flow diagram included (14). | Confirm complete PRISMA-compliant diagram is present. |  |
| Database search strategy documented (8.1). | Verify databases, date ranges, exact search terms, and filters are specified. |  |
| Record counts match PRISMA phases (14 vs 8.1). | Cross-check counts across sections for consistency. |  |
| Inclusion/exclusion criteria match PICO (8.1.1). | Ensure criteria explicitly reference PICO elements. |  |
| Full retrieval list provided (8.1.1, 13). | Verify all articles (included and excluded) have complete citations. |  |
| Unique IDs assigned to each article (8.1.1). | Check IDs (e.g., PUB-001, EMB-001) are applied consistently. |  |
| Inclusion/exclusion table complete (8.1.1). | Confirm rationale linked to article IDs. |  |
| Included study appraisal (8.3). | Verify appraisal links to corresponding article IDs. |  |
| SOTA benchmarks defined (5.5). | Confirm benchmarks are based on peer-reviewed literature. |  |
| Device outcomes mapped to SOTA (9.2.2). | Verify safety/performance mapping table exists. |  |
| Equivalent devices referenced (8.4). | Confirm cited in mapping tables and compared. |  |
| PMS data incorporated (9.2.2). | Verify PMS outcomes are integrated into SOTA mapping. |  |
| Conclusions align with mapping (10.2.2, 11.1). | Check final conclusions match benchmark analysis. |  |

Approval

Author: ___________________		Date: ___________

Reviewer: ___________________		Date: ___________

Approval: ________________		Date: ___________