Tulip Medical

Quality System Procedure

QSP 9.1-5

Advisory Notices, Recalls & Vigilance, Medical Device Incident Reporting & Field Safety Corrective Actions

Approvals

President								Date

V.P. Operations	 						Date

V.P. Finance							Date

V.P. Regulatory Affairs/Management Rep.				Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 1 | N/A | Types of Reporting added. |
| 2 | 09/23/2013 | Medical Device Problem Reporting added. |
| 3 | 09/08/2014 | Updated references to ISO 13485:2012. |
| 4 | 07/17/2015 | Complete rewrite for EU and Canadian recalls. |
| 5 | 09/11/2018 | Reformatted document |
| 6 | 10/31/2018 | Document title correction from “Recalls, Corrections and Removal” back to it’s original Document title: “ Advisory Notices, Recalls & Vigilance. |
| 7 | 11/06/2018 | Revised sections 1.0, 3.0, 4.0, 7.2.1, 7.2.2, 7.2.4, 7.2.2.5, 7.3.2.1, 7.3.2.2, 7.3.3.2, and 7.4.1.1 to read “Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan” or reference it. |
| 8 | 10/28/2019 | Section 2.0 revised the existing erroneous Purpose & Scope. Section 3.0 updated all reference documents. Section 3.0 added market-specific reporting guidance chart. 7.0 Added requirements of US, EU, CANADA, JAPAN, BRAZIL, and AUSTRALIA |
| 9 | 1/13/2022 | Added Medical Device Incident Reporting & Field Safety Corrective Actions to Title – Rewrote Purpose-Scope - Added EU 2017/745 and MIR in Reference section – Expanded Definitions – Expanded section 4-Abbreviations and section 5-Responsibilities - Added sections 7.2, 7.6, 7.7, 7.8, 7.9, 7.10, 7.11 – Added reference to QSP 8.2-2 Customer Feedback and Complaints - Added Sections 8 and 9 – Added Appendix A and B |
| 10 | 07/25/2023 | Updated Address of EC Representative MedEnvoy Europe in Section 7.11.B.1 Updated EC Address per Quality Plan QP 016.4. |
| 11 | 7/14/2025 | Added Section 7.1 “Responsibilities and Action Flow for Incident Reporting” to clarify MDR reporting roles, responsibilities, and workflows. Integrated direct references to MDR Determination Forms (8.2-2-4 to 8.2-2-7) and the complaint process. Added Appendix A: Responsibility Flow Matrix. Update AR from Emergo to MedEnvoy including, emails and address through out the document. |
| 12 | 08/06/2025 | Section 2.0 Replaced broken hyperlinks |

- PURPOSE -- SCOPE

The purpose of this procedure is to provide methods and forms used for reporting serious incidents and to initiate a field safety corrective action (FSCA), or product retrieval of any product manufactured and/or distributed by Tulip Medical Products to the appropriate Competent Authority (CA) in the EU and per the contract with EU Authorized Representative (AR).

These procedures apply to all situations which could conceivably result in a recall or product retrieval action. It also describes the procedures for reporting adverse events for products distributed in the U.S., Europe, Canada Brazil, Japan and Australia.

This procedure encompasses the requirements for reporting and record-keeping requirements for information received, which reasonably suggests that a manufactured device may have caused, or contributed to, a death, serious injury or has malfunctioned, and that the device would likely cause, or contribute to, a death or serious injury, if the malfunction were to recur.

This procedure also applies to all situations that could conceivably result in a FSCA

- REFERENCE DOCUMENTS

| ISO 13485:2016 Medical Devices – 8.3.3 Actions in response to nonconforming product detected after delivery EU 2017/745 Medical Device Regulation Manufacturer Incident Report (MIR) for Serious Incidents (MDR / IVDR) (effective January 2020): https://ec.europa.eu/docsroom/documents/37348 |  |
| --- | --- |
| US FDA 21 CFR Part 806 – https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-806 US FDA 21 CFR Part 7 – https://www.ecfr.gov/current/title-21/chapter-I/subchapter-A/part-7 MEDDEV 2.12 - https://ec.europa.eu/docsroom/documents/32305/attachments/1/translations Form 9.1-5-1 Advisory Notices, Recalls & Vigilance Plan Health Canada SOR/98/282 |  |
| https://www.canada.ca/en/health-canada.html |  |
| https://recalls-rappels.canada.ca/en |  |
| Form 9.1-5-1 | Advisory Notices, Recalls & Vigilance Plan |
| QSP 4.2-1 | Document Control |
| QSP 8.3-1 | Control of Nonconforming Product |
| QSP 6.2-1 | Training and Qualification |
| QSP 9.1-2 | Conducting Recalls – US FDA |
| QSP 9.1-6 | Medical Device Reporting |
| QSP 9.1-7 | Post Market Surveillance |
| QSP 8.2-2 | Customer Feedback and Complaints |

| REPORTS TO REGULATORY AUTHORITIES |  | When a medical device malfunctions and contributes to serious patient injury or death, the governing laws in the country where the incident occurred are applicable to reporting requirements of the device manufacturers, users and distributors. |
| --- | --- | --- |
| Guidance and procedures pertaining to Adverse Events |  |  |
| Market | Regulation Document | Title |
| EU | EN ISO 13485:2016 | Medical Device Regulation http://www.isomcx.com/uploadfile/2018/0623/20180623095340510.pdf |
| EU | EU 2017/745 (MDR) | EU Medical Device Regulation https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ:L:2017:117:TOC https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=OJ:L:2017:117:FULL |
| EU | MDD 93/42 EEC | EU Medical Device Directive https://ec.europa.eu/growth/single-market/european-standards/harmonised-standards/medical-devices_en |
| EU | MEDDEV 2.12/1 | https://ec.europa.eu/docsroom/documents/32305/attachments/1/translations |
| USA | FDA – How to Report Medical Device Problems | https://www.fda.gov/medical-devices/medical-device-safety/medical-device-reporting-mdr-how-report-medical-device-problems |
| USA | FDA – Recalls, Corrections and Removals | https://www.fda.gov/medical-devices/postmarket-requirements-devices/recalls-corrections-and-removals-devices |
| USA | FDA 21 CFR 806 | US FDA PART 806—Medical Devices; Reports of Corrections and Removals https://www.ecfr.gov/current/title-21/chapter-I/subchapter-H/part-806 |
| USA | FDA 21 CFR | US FDA Code of Federal Regulations https://www.ecfr.gov/current/title-21 |
| CANADA | SOR-98-282 | Health Canada’s Medical Device Regulations (CDMR) https://laws-lois.justice.gc.ca/eng/regulations/sor-98-282/ |
| BRAZIL | ANVISA RDC No. 67/2009, Section II | Brazil’s Adverse Events https://www.fda.gov/media/93179/download |
| AUSTRALIA | TGA | TGA’s Adverse Events https://www.tga.gov.au/reporting-adverse-events |
| JAPAN | PMDA (Pharmaceuticals and Medical Devices Agency) | JAPANs PMDA https://www.pmda.go.jp/english/about-pmda/0004.html |

- DEFINITIONS

Abnormal Use: act or omission of an act by the operator or user of a medical device as a result of conduct which is beyond any means of risk control by the manufacturer

Authorized representative (Art. 2(32)): any natural or legal person established within the Union who has received and accepted a written mandate from a manufacturer, located outside the Union, to act on the manufacturer's behalf in relation to specified tasks with regard to the latter's obligations under the MDR.

Caused or Contributed to: a device may have “caused or contributed to” a death or serious injury, if the death or serious injury was or may be attributed to the device, or the device may have been a factor in the death or serious injury because of device failure or malfunction, improper or inadequate design, manufacture, labeling, instruction or user error.

Complaint: any communication about a device that alleges deficiencies related to the identity, quality, durability, safety, effectiveness, or performance after it is released for distribution.

Corrective Action (MDR Article 2(67)): action taken to eliminate the cause of a potential or actual non-conformity or other undesirable situation.

Device Deficiency (MDR Article 2(59)): any inadequacy in the identity, quality, durability, reliability, safety or performance of an investigational device, including malfunction, use errors or inadequacy in information supplied by the manufacturer.

Eudamed: the European database on medical devices that integrates different electronic systems to collate and process information regarding devices on the market and the relevant economic operators, certain aspects of conformity assessment, notified bodies, certificates, clinical investigations, vigilance and market surveillance. {Currently delayed until May 2022}

European Union (EU): collectively referred to as “EU” and includes 27 official EU member states, the EEA (Norway, Iceland, and Liechtenstein), and Switzerland.

Field Safety Corrective Action (FSCA)-EU (MDR Article 2(69)):  – Corrective action taken by a Manufacturer for technical or medical reasons to prevent or reduce the risk of a serious incident in relation to a device made available on the market. Such actions, whether associated with direct or indirect harm, should be reported and should be notified via a Field Safety Notice. The FSCA may include:

- The return of a medical device to the supplier;

- Device modification;

- Device exchange;

- Device destruction;

- Retrofit by purchaser of manufacturer’s modification or design change;

- Advice given by manufacturer regarding the use of the device and/or the follow up of patients, users or others (e.g. where the device is no longer on the market or has been withdrawn but could still possibly be in use e.g. implants or change in analytical sensitivity or specificity for diagnostic devices).

Incident (MDR Article 2(64)): any malfunction or deterioration in the characteristics or performance of a device made available on the market, including use-error due to ergonomic features, as well as any inadequacy in the information supplied by the manufacturer and any undesirable side-effect.

Intended Purpose (MDR Article 2(12)): the use for which a device is intended according to the data supplied by the manufacturer on the label, in the instructions for use or in promotional or sales materials or statements and as specified by the manufacturer in the clinical evaluation.

Making Available on the Market (MDR Article 2(27)): any supply of a device, other than an investigational device, for distribution, consumption or use on the Union market in the course of a commercial activity, whether in return for payment or free of charge.

Malfunction: failure of a device to meet its performance specifications or to perform as intended.  A malfunction is reportable when it is likely to cause or contribute to a death or a serious deterioration in the state of health or serious injury of a patient, user or other person.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (QSPs).

Manufacturer (MDR Article 2(30)): the natural or legal person who manufactures or fully refurbishes a device or has a device designed, manufactured or fully refurbished, and markets that device under its name or trademark.

Manufacture Incident Report (MIR): the MDR requires manufacturers to report incidents where death or serious deterioration in the state of health of a person have occurred or near incidents which might have led to the death or serious deterioration in the health. Refer to MDR Art. 87 through 92.

Market Withdrawal – Removal or correction of a distributed product, which involves a minor violation that would not be subject to legal action by the FDA or which involves no violation. (MDR Article 2(63) any measure aimed at preventing a device in the supply chain from being further made available on the market.

Removal – The physical removal of a device from its point of use to some other location for repair, modification, adjustment, relabeling, destruction or inspection.

Reporting to regulatory authorities -  If applicable regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the organization shall document procedures for providing notification to the appropriate regulatory authorities.

Risk to Health – A reasonable probability that use of, or exposure tom the product will cause serious adverse health consequences, or death; or, that use of, or exposure to, the product may cause temporary or medically reversible adverse health consequences, or an outcome where the probability of serious adverse health consequences is remote.

Serious Incident (MDR Article 2(65)): any incident that directly or indirectly led, might have led or might lead to any of the following:

    - the death of a patient, user or other person,

    - the temporary or permanent serious deterioration of a patient’s, user’s or other person’s state of health,

    - a serious public health threat.

    - Serious deterioration in state of health -

- Life threatening illness / injury;

- Injuries or illness that result in permanent damage to a body structure or permanent impairment of a body function;

- A condition necessitating medical or surgical to prevent the two (2) bullet points above;

- Any indirect harm as a consequence of an incorrect diagnosis or IVD test results when used within the manufacturer instruction for use; and

- Fetal distress, fetal death or congenital abnormality or birth defects.

Serious Public Health Threat (MDR Article 2(66)): an event which could result in imminent risk of death, serious deterioration in a person’s state of health, or serious illness, that may require prompt remedial action, and that may cause significant morbidity or mortality in humans, or that is unusual or unexpected for the give place and time.

Stock Recovery – The correction or removal of a device that has not been marketed or that has not left the direct control of the manufacturer, i.e. the device is located on the premises owned, or under the control of, the manufacturer, and no portion of the lot, model, code or other relevant unit involved in the corrective or removal action has been released for sale or use.

Technical Documentation (TD): the name given to the EU regulatory documentation / file for all medical devices.

Recall – See below.

US: a firm’s removal or correction of a marketed product that the FDA considers to be in violation of the laws it administers and against which the agency would initiate legal action, e.g. seizure. Recall does not include a market withdrawal or a stock recovery.  See QSP 9.1-2 Recalls – US FDA. https://www.fda.gov/medical-devices/postmarket-requirements-devices/recalls-corrections-and-removals-devices

Canada: in respect of a medical device that has been sold, means any action taken by the manufacturer, importer or distributor of the device to recall or correct the device, or to notify its owners and users of its defectiveness or potential defectiveness, after becoming aware that the device Health Canada’s Medical Device Regulations (CDMR) https://www.canada.ca/en/health-canada/services/drugs-health-products/medical-devices/legislation-guidelines.html

Europe: see definition of Field Safety Corrective Action (FSCA)-EU – Section 8.6  (MDR Article 2(68)) https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=OJ:L:2017:117:TOC

Japan: JAPANs PMDA see https://www.pmda.go.jp/english/about-pmda/0004.html

Brazil: see https://www.fda.gov/media/93179/download

Australia:  see TGA’s Adverse Events https://www.tga.gov.au/reporting-adverse-events

- ABBREVIATIONS

AR – Authorized Representative

CA – Competent Authority

EEA – European Economic Area

EFTA – European Free Trade Association

EU – European Union

FSCA – Field Safety Corrective Action

MDR – Medical Devices Regulation

MIR – Manufacturer Incident Report

NCR – Nonconformance Report

PRRC – Person Responsible for Regulatory Compliance

TD – Technical Documentation

MR – Management Representative

QA – Quality Affairs

- RESPONSIBILITIES

- The PRRC or MR is responsible for ensuring compliance with this procedure and assuring the effectiveness for any serious incidents, Manufacturer Incident Reports (MIR), and FSCAs, and for reporting to our EU AR and applicable CAs.

- The PRRC will also perform a review, as necessary, to determine whether an MIR and/or FSCA should be initiated and will provide appropriate corrective action as necessary.

- In this review, the PRRC will ascertain all available facts and review with the appropriate personnel to determine whether vigilance report should be initiated and assure that a strategy and procedures are carried out in an orderly fashion.

- The PRRC will perform a review as necessary to determine whether a MIR shall be filed, based on the determination of Reportable Event Determination checklist.  Serious incidents will be reviewed with the Executive Management Team and other applicable personnel.

- GUIDANCE

What is an Adverse Event?  An adverse event refers to any unfavorable occurrence involving the device(s).  However, the event must be reported to the FDA if the device has:

- Caused or contributed to a death or serious injury.

- Malfunctioned and could potentially cause or contribute to a death or serious injury if the malfunction were to occur.

What is a device recall?  Tulip can elect to recall any device when there is a risk of death or serious deterioration to the state of health caused by a device.  A recall is the removal or correction of a marketed product that the FDA considers to be in violation of the laws it administers.

QSP 9.1-2 Conducting Recalls – The US FDA defines and describes how to conduct product recalls in accordance with the US regulations.

FDA 21 CFR Part 803 specifies the types of reports and data that must be included in each eMDR submission.

Table 1

| Regulation | Requirement | Report Time | Reporting to |
| --- | --- | --- | --- |
| FDA 21 CFR Part 803 | Manufacturers are required to report to the FDA when they learn that any of their devices may have caused or contributed to a death or serious injury | Within 30 calendar days of becoming aware of an event. | FDA. Required Form FDA 3500A; |
| FDA 21 CFR Part 803 | 5-Day reports for an event designated by the FDA or an event that requires remedial action to prevent an unreasonable risk of substantial harm to the public health | Within 5 calendar days of becoming aware of an event | FDA Required Form FDA 3500A; |
| FDA 21 CFR Part 803 | User Facilities must report all device-related deaths or serious injury | Within 10 workdays of becoming aware. | FDA & Manufacturer |

For the EU, an incident report must be filed if a device malfunction, deterioration in performance, inadequate instructions, or inadequate labeling resulted in death or serious injury or may lead to death or serious deterioration in state of health if it were to reoccur.

Table 2

| Regulation | Requirement | Report Time | Report to |
| --- | --- | --- | --- |
| MEDDEV 2.21/1 | Serious public health threat | 2 days | FSCA Reports to Competent Authority or the member state where the incident(s) occurred. |
| MEDDEV 2.21/1 | Death or serious deterioration of health | 10 days | FSCA Reports Competent Authority or the member state where the incident(s) occurred. |
| MEDDEV 2.21/1 | Serious incidents | 15 days | FSCA Reports Competent Authority or the member state where the incident(s) occurred. |

- PROCEDURE

Different markets around the world have different reporting requirements. Follow the table below for the reporting requirements of the US, Japan, the EU, Canada, Brazil, and Australia.

Table 3

| Market | Severity | Report within | Who Reports | Annual Report Required | Reporting individual incidents occurring in other countries? |
| --- | --- | --- | --- | --- | --- |
| US | Death, Serious Injuries, or Malfunctions | 30 days | Manufacturer, User | Yes | Yes within 30 days |
|  | Remedial Action Needed to Prevent Risk of Substantial Harm | 5 days | Manufacturer, User | Yes | Yes*** within 30 days |
| JAPAN | Death plus malfunctions, faults or breakage** | 15 days | Market authorization holder | No | No |
|  | Serious events | 30 days |  |  |  |
| EU * | Serious public health threat | 2 days | Manufacturer | No | No only applies to FSCA |
|  | Death or serious deterioration of health | 10 days |  |  |  |
|  | Serious incidents | 15 days |  |  |  |
| CANADA | Death or serious deterioration of health | 10 days | Manufacturer & Importer | No | No |
|  | Recurrence could lead to death or serious health deterioration | 30 days |  |  |  |
| BRAZIL | Death or serious deterioration of health | 72 hours | Manufacturer & Registration Holder | No | Yes*** within 10 days |
|  | Serious adverse event but no deaths | 10 days |  |  |  |
|  | Reoccurrence could lead to death or serious health deterioration | 30 days |  |  |  |
| AUSTRALIA | Serious public health threat | 48 hours | Manufacturer & Sponsor | Yes | No |
|  | Death or serious deterioration of health | 10 days |  |  |  |
|  | Reoccurrence could lead to death or serious health deterioration | 30 days |  |  |  |
| * European requirements are for the Medical Device Regulation (2017/745). ** Japan has specific reporting timelines for incidents involving device malfunctions, breakages, and fault that could lead to serious events. *** Only required if you also sell the same device in this market. Types of reportable events vary by country. The US FDA requires the reporting of all ex-US events but in Brazil, for example, this only applies to “deaths, serious deterioration of health, and counterfeit devices.” |  |  |  |  |  |

Note: For adverse events resulting from Customer Feedback or Complaints see QSP 8.2-2 Customer Feedback and Complaints

  - Advisory Notices, Recall and Vigilance.

This section defines the roles and sequence of actions for identifying, evaluating, and reporting serious incidents, adverse events, and field safety corrective actions (FSCAs) under MDR 2017/745 Articles 87–92, as well as other applicable global vigilance requirements.

Tulip Medical Products uses a complaint-integrated MDR reporting process. All complaints are initially captured using one of the Customer or Distributor Complaint Forms (Forms 8.2-2-0 through 8.2-2-2). The Complaint Investigation Form (8.2-2-3) includes a mandatory MDR Determination, which must be completed within 48 hours of receiving the completed questionnaire.

Regional MDR Determination Forms include:
- Form 8.2-2-4 – MDR Determination – FDA
- Form 8.2-2-5 – MDR Determination – Canada and Australia
- Form 8.2-2-6 – MDR Determination – Brazil and Japan
- Form 8.2-2-7 – MDR Determination – EU

The determination of whether an incident is reportable and which regulatory authority must be notified is the responsibility of the Person Responsible for Regulatory Compliance (PRRC). The PRRC may consult with RA staff and Quality as needed, but retains final accountability for compliance with reporting obligations.

Timelines for submission are based on incident type, as defined in QSP 9.1-5 Section 7.7 and Appendix A (EU reporting: 2-day, 10-day, 15-day rule). All timelines begin the day after Tulip becomes aware of the event.

Completed MDR reports are submitted either directly by Tulip or, in the EU, via the EU Authorized Representative (currently MedEnvoy). All records are filed in accordance with QSP 4.2-1 and cross-referenced with the associated complaint file and any resulting CAPA.

Responsibility Flow Matrix

| Step | Action/Decision | Responsible Party | Reference Form(s) | Timeline |
| --- | --- | --- | --- | --- |
| 1 | Intake of complaint | Customer Service / Distributor | Forms 8.2-2-0 to 8.2-2-2 | Within 1 business day of receipt |
| 2 | Entry into complaint system | DCA / RA | Form 8.2-2-3 | Same day |
| 3 | MDR Determination | PRRC (in consultation with RA) | Forms 8.2-2-4 to 8.2-2-7 | Within 48 hours of completed complaint |
| 4 | Finalize report content | Regulatory Affairs | Region-specific MIR or FSCA templates | Prior to submission deadline |
| 5 | Submit report | PRRC or EU AR (e.g., MedEnvoy) | Email or EUDAMED (EU) | 2 / 10 / 15 calendar days depending on severity |
| 6 | File and archive | Document Control Administrator | Complaint File + QMS | Upon report submission |
| 7 | Initiate CAPA if needed | QA / RA / PRRC | CP Forms | Within 30 days |
| 8 | Track effectiveness | Management Representative | QMS Audit Schedule | Per internal audit cycle |

  - For EU, go to section 7.7; all others continue below.

  - For Canada

  - PRRC agrees to follow Advisory Notices, Recalls, and Vigilance Plan

  - The QA Manager

    - Prepares Form 9.1-5-1 Advisory Notices, Recalls, and Vigilance Plan

    - Ensures completion and review of the Advisory Notices, Recalls, and Vigilance Plan. The QA Manager gathers all necessary information regarding the issue, i.e., issue description, product assessment, quantity involved, lot number(s) involved, etc., and determines activities (e.g., correction or removal, repair, modification, adjustment, relabeling, destruction, and/or inspection).

    - Determine whether Reportable or not

- Reportable – Document activities and submit report(s) to appropriate legal agency(cies) within the prescribed timeframes, as noted in this procedure.

- Not Reportable – Document activities and report internally.

    - Determine the depth of the Recall and document on Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan the appropriate level in the distribution chain that the recall is to extend (e.g., general public, retail, user, distributor).

    - Complete Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan This includes documenting the activities related and attaching any reference documents, such as protocols, procedures, etc., that provide support and background to the activity.

  - Top Management

    - Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan must be approved by Top Management.

    - Is customer communication required?

      - If YES, prepare/approve Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan

      - If NO, document the decision and continue with Advisory Notices, Recalls and Vigilance Plan activity.

      - The following key elements must be included in the customer communication associated with a C&R and/or Advisory Notice.

- Description of the issue

- A listing of the product affected [product name, code, size, lot number(s)/serial number(s)]

- The associated safety risk

- Actions taken by Tulip Medical to address the issue

- Actions the customer will need to take to mitigate the potential safety risk

- Contact information (day of the week, time of day and time zone contact available)

- A customer reply form, if applicable

- A statement regarding regulatory body notification, if applicable

    - Perform due diligence regarding customer notification. Check effectiveness.

      - Due diligence will be performed by each affected region as defined per regional/local regulations and any additional procedures documented on Form 9.1-5-1 Corrections and Removals Assessment and Action Plan. At a minimum, the following steps shall be taken:

- Undeliverable communications will be investigated to verify the correct address was used or if the customer is still valid.

- At least three (3) attempts must be made after the initial notification to get confirmation of successful delivery in a timely manner.

- Additional attempts may be required per the direction of the Top Management, or the regulatory agency/MoH.

- Effectiveness of the customer communication will be monitored monthly.

      - Objective evidence or successful delivery of the communication will be obtained via one or a combination of the following methods:

- Completed Customer Reply Form.

- Signed Receipt of Mail Delivery.

- Documentation of Verbal Confirmation (include name/title of individual contacted, the date/time of conversation, and name of individual confirming receipt).

- Other, as documented on Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan

  - QA manager

      - Initiate close and document on Form 9.1-5-1 Advisory Notices, Recalls and Vigilance Plan Based on regional/local requirements, the Authorized Representative will indicate their closure of the Advisory Notices, Recalls and Vigilance Plan in writing to the Top Management Representative. This will be based on the following criteria:

- Due diligence has been demonstrated that all possible affected customer responses have been received.

- Due diligence has been demonstrated that all possible affected product has been recovered, corrected, reconditioned or destroyed.

- Reconciliation of all product holds.

- Receipt of documentation of the completion of any associated correction.

- Based on regional/local requirements, the regulatory body(s)/MoH(s) closure documentation associated with the country(s) affected by the Advisory Notices, Recalls and Vigilance Plan has been received.

      - Once the PRRC receives notification from all affected regions and verifies that all of the Correction & Removals file requirements have been met per this procedure, the PRRC will close the action.

  - Competent Authority (CA) (Art. 87)

    - Where a CA obtains such reports on suspected serious incidents from healthcare professionals, users or patients, the CA will ensure that the manufacturer of the device is informed of the suspected serious incident without delay.

      - If PRRC determines that the incident meets the definition of a serious incident, the manufacturer will provide a MIR report to the CA of the member state in which the serious incident occurred and shall take the appropriate follow-up action.

      - If PRRC considers the incident is not a serious incident or is an expected undesirable side-effect, PRRC will provide an explanatory statement to the CA. If the CA does not agree with the conclusion of the explanatory statement, the CA may require a MIR.

    - CAs may conduct their own assessments on the trend reports and require Tulip Medical Products to 	adopt appropriate measures in order to ensure the protection of public health and patient safety.

  - Utilize the data in Table 3 to determine reporting time and type of report required for each type of event that has a reportable incident.

  - Where necessary to ensure timely reporting, submit an initial report (refer to MIR Form) that is incomplete followed up by a complete report (refer to MIR Form).

  - Medical Incident Reporting

    - Use Table 3 or Appendix A and Art. 87 to determine reporting time and type of report required.

    - Serious Incident Reports are submitted using Eudamed (Art. 92) using the MIR Form to the CA where the incident occurred.

      - MedEnvoy can submit on behalf of Tulip Medical Products. If MedEnvoy is submitting the report, the file shall be sent to MedEnvoy at:

EU-AR@medenvoyglobal.com

MedEnvoy Global B.V.

Prinses Margrietplantsoen 33 – Suite 123

2595 AM The Hague

The Netherlands Phone: +31.70.326.2148

    - The complete list of CAs for the EU is found on the web site:    http://ec.europa.eu/growth/sectors/medical-devices/contacts/index_en.htm.

    - If a CA verbally requests additional information, PRRC shall ask to follow up such requests in writing.

    - A separate file must be maintained for MIRs.

  - Analysis of Serious Incidents (Art. 89)

    - Following a report of a serious incident, Tulip Medical Products must perform the necessary investigation in relation to the serious incident and the devices concerned.

      - This investigation shall include a risk assessment of the incident.

      - No action can be performed during the investigation which involves altering the device or a sample of the batch concerned in a way which may affect any subsequent evaluation of the causes of the incident, prior to informing the CA of such action.

    - At the conclusion of the investigation, a Final Report (refer to MIR Form) must be submitted to the CA setting out Tulip Medical Products’s findings and relevant corrective actions to be taken, if appropriate.

    - Appendix B provides an overview of the steps that will be taken in response to the MIR reporting.

  - Trend Reports (Art. 88)

    - Tulip Medical Products must report any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side-effects that could have a significant impact on the benefit-risk analysis and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits.

    - Tulip Medical Products must establish a method to manage the incidents referred to in 7.11 and the methodology used for determining any statistically significant increase in the frequency or severity of such incidents, including the observation period in the Post-Market Surveillance Plan, refer to QSP 9.1-7 Post Market Surveillance and Form.

- Handling Field Safety Corrective Action (FSCA)

  - Should the evidence point to the need for a FSCA and/or market withdrawal, specific reporting action may be required by all CAs.

Notify the AR, MedEnvoy , via email at EU-AR@medenvoyglobal.com

    - if the products are sold or marketed in the EU for notification of the FSCA.

  - Reporting to Competent Authorities (Cas)

    - Except in cases of urgency in which Tulip Medical Products needs to undertake FSCA immediately, Tulip Medical Products shall, without undue delay, report the FSCA in advance of the FSCA being undertaken.

    - Submit the FSCA Report to the appropriate CA via EUDAMED (or via email if EUDAMED is not operational).

If Tulip Medical Products has mandated that MedEnvoy will submit reports on Tulip Medical Product’s behalf. The report form and FSN shall be sent via email to MedEnvoy at EU-AR@medenvoyglobal.com

    - The FSCA is provided on the European Commission website: http://ec.europa.eu/growth/sectors/medical-devices/guidance/index_en.htm.

    - A template of the FSN that can be used to notify customers is provided on the European Commission website.

    - The AR can assist in the execution of the FSCA in the EU.

    - If a CA verbally requests additional information, PRRC shall ask to follow up such requests in writing.

    - Tulip Medical Products will notify each importer / distributor / customer and maintain records of each attempt (date, method, result).

    - Response forms shall be compiled and used for reconciliation of data.

    - Closure of a FSCA shall be based on the submittal of a Final FSCA report detailing the reconciliation of distributed product in Europe and closure of CAPA actions and/or receipt of a closeout report from a Competent Authority. Note that not all CAs will provide close-out reports.

    - Each CA which received an Initial FSCA report, must receive a Final FSCA report.

    - A separate file must be maintained for all FSCAs.

- QUALITY RECORDS

  - The following information must be retained in the appropriate file.  Each are maintained in a separate file by market location and/or cross-referenced as needed within the QMS.

    - A copy of the original complaint record and/or non-conformance report.

    - Documentation of follow-up and attempts to follow-up and obtain additional information about the event. If applicable, an explanation of why certain information cannot be obtained.

    - A copy of the Health Hazard Evaluation, for FSCAs.

    - Copies of any relevant test reports and investigation reports.

    - Documentation related to the decisions used to determine the reportability of the event.

    - Documentation of the final assessment of the event and any corrective or remedial action taken.

    - Copies of reports and other information submitted to the applicable CAs.

    - Copies of information submitted to AR, importers, distributors and / or customers.

    - Copies of any other information required regulations.

  - Closure of a MIR or serious incident shall be based on the results of the investigation and/or receipt of a closeout report from a CA.

  - Closure of FSCA shall be based on the completion of all corrective actions and/or receipt of a closeout report from a CA.

    - If remedial action is taken, and includes a significant change or modification to the device, a determination shall be made by PRRC or other designated personnel, as to whether or not an update is required to product TD and product registration (inform AR, as applicable).

  - Retention times for MIR and FSCA files must be in accordance with QSP 4.2-1.

Appendix A: Reporting Types and Timelines

| # If device is… CE marked device | MIR Action |  |
| --- | --- | --- |
|  | Incident occurred in EU | Incident occurred outside EU |
| Reporting Timeframe for 2, 10, & 15-day reports, start the day after the first Company employee is notified. Report is to be sent within the time specified below. |  |  |
| Serious Public Health Threat | Yes Initial Report immediately but not more than 2 calendar days | No* |
| Death or unanticipated serious deterioration of health has occurred and was caused or contributed to by a device manufactured by our company. | Yes Initial Report immediately but not more than 10 calendar days | No* |
| All other serious incidents | Yes Initial Report immediately but not more than 15 calendar days | No* |
| Additional information becomes available after the initial report has been submitted. | Final Report | N/A |
| A FSCA has been initiated due to risk of Death or Serious Injury (reported to a Regulatory Agency). | Yes, via FSCA report if any affected devices are marketed in the EU* On or before initiating FSCAs Refer to QSP 9.2-3 | No, if any affected devices are marketed in the EU* |

*Incidents which occurred outside the EU areas should be considered for reporting with recommendation to generate a rationale or justification documenting the reason for not notifying or advising other regulatory agencies.  This is particular true if the incident did not occur in these regulatory regions, however, the product or product family is marketed or sold in these regulatory regions.  Generally, if the product or product family is not marketed or sold in these regulatory regions, then notification is not required.

Appendix B: Medical Device Incident Reporting Sequence

The United Kingdom left the European Union on 31 Jan. 2020 (BREXIT); the future regulatory situation is currently unclear.

Until Eudamed is functioning, reports are submitted via email to the Vigilance Contact Point listed on the EU Commission website.