Tulip Medical

Quality System Procedure

QSP 7.3-1

Design & Development

Approvals

VP Regulatory Affairs	Date

Regulatory Affairs Coordinator	Date

Document Control Administrator	Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 04/23/2014 | Released documentation. |
| 1 | 09/22/2014 | Updated ISO references. |
| 2 | 09/23/2014 | Updated ISO references to 13485:2016. |
| 3 | 08/30/2018 | Updated sections to meet requirements of ISO 13485:2016. |
| 4 | 05/07/2020 | Added requirements for Essential Principles Checklist (7.3.1.1); added reference to state of the art standards (7.3.1.2); changed font to Calibri and size to 10 |
| 5 | 07/08/2020 | Added reference to Australian standard requiring essential principles checklist (7.3.1.1) |
| 6 | 3/14/2024 | Added note to section 7.3 to add safety & efficacy consideration Added design requirements from SOR-98-282 to section 7.6.6 |
| 7 | 5/23/2025 | Updated all references from MDD to MDR in accordance with Regulation (EU) 2017/745. Included appropriate citations to Annex I (GSPR) and Annex XV (Clinical Investigations). Added safety and efficacy requirements in design input section to align with MDR principles. |
| 8 | 07/12/2025 | 7.10.1 Change Request Process — Updated to clarify the requirement for documenting reasons for change, with reference to QSP 4.2-3 and examples (e.g., supplier issues, audits, test deviations). |

- PURPOSE

To detail the process for control of the design and development for new products or significant changes to medical devices developed by Tulip Medical.

Note: This procedure first became effective in 2013. Any products that Tulip Medical owns the design and of and introduced into the market before 2013 will not have had a design process which followed the requirements of this procedure exactly.  Instead, for such products we have collected records and have retroactively created a Design History File to the best of our ability. For such products:

- We have Design History Logs. The Design History Logs take the place of Design and Development Planning for such products. The Design History Logs note on the Log, “no design control procedure was in place at the time of the design and development of this product. We have recreated the design history as best as possible.”

- The Design History Logs indicate when the product was first introduced into the market. The Design History Logs also show when the product was accepted into production, i.e., when the Design Transfer took place.

- We have a Design History File summary using existing records. The Design History File Summary for products introduced into the marketplace before this procedure became effective may not be as complete as products developed after this procedure became effective.

From the first effective date of this QSP on, Tulip Medical will follow the requirements of this QSP.

- SCOPE

This Quality System Procedure is to be followed by all Tulip Medical Products personnel. Comprehension of, and satisfactory expertise in, the material covered by this Quality System Procedure by each employee will be indicated through the completion of all applicable and appropriate training requirements, as is indicated by the completion of Form 6.2-1-2 Training Record in accordance with QSP 6.2-1 Training and Qualification.

- REFERENCE DOCUMENTS

| Food and Drug Administration 21 CFR 820 Quality System Regulation |  |
| --- | --- |
| International Standard Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes ISO 13485:2016 |  |
| Regulation (EU) 2017/745 on Medical Devices (MDR) |  |
| SOR/98-282 Health Canada Medical Device Regulations |  |
| BS EN ISO 14155-1:2011 Clinical Investigation of medical devices for human subjects – Part 1: General Requirements |  |
| International Standard Medical devices – Application of risk management to medical devices ISO 14971:2012 Medical Devices Essential Principles Checklist |  |
| Form 4.2-1-2 | Change Request |
| Form 6.2-1-2 | Training Record |
| Form 7.3-1-1 | Design Plan |
| Form 7.3-1-2 | Design Inputs / Design Outputs |
| Form 7.3-1-3 | Design Review |
| Form 7.3-1-4 | Design Verification / Validation Summary |
| Form 7.3-1-5 | Design Transfer Checklist |
| QSP 4.2-1 | Document Control |
| QSP 4.2-2 | Control of Quality Records |
| QSP 7.3-2 | Device Master Record |
| QSP 7.3-3 | Risk Management |

- DEFINITIONS

Controlled Documents – Quality System documentation, including the Quality Manual, ASPs, Work Instructions, forms, product specifications, product drawings, Device Master Records, product labeling, downloadable instructions, and external standards and regulations.

A document is defined as controlled and considered current if, and only if, it is located in the ‘Controlled Documents’ section of the company’s computer server. There is one, and only one, copy of each controlled document.

Design and Development Planning – Plans that describe or reference the design and development activities and define responsibility for implementation.

Design History File – A compilation of records which describes the design history of a finished device.

Design Inputs – The physical and performance requirements of a device that are used as a basis for the device design.

Design Outputs – The results of a design effort at each design phase and at the end of the total design effort. The finished design output is the basis for the initial Device Master Record. The total finished design output consists of the device, its packaging and labeling and the Device Master Record. The design outputs are intended to meet the requirement of the design inputs.

Design Review – A documented, comprehensive, systematic examination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements and to identify problems.

Design Transfer – The procedure by which the device design is correctly translated into production specifications.

Device Master Record – A compilation of the procedures and specifications for a finished device.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (ASPs).

Process Validation: Confirmation by objective evidence that a process consistently produces the intended result.

State of the Art – developed stage of technical capability at a given time as regards products, processes, and services, based on the relevant consolidated findings of science, technology and experience.

Verification – Confirmation through the provision of objective evidence that specified requirements have been fulfilled.

Note: Confirmation can include activities such as performing alternative calculations, comparing a new design with a similar proven design, undertaking tests and demonstrations and reviewing documents.

Validation – Confirmation through the provision of objective evidence that the requirements for a specific intended use or application have been fulfilled.

Note: The use conditions can be real or simulated.

- ABBREVIATIONS

DHF – Design History File

MR – Management Representative

- RESPONSIBILITIES

The Management Representative is responsible for maintaining this procedure in a current state.

The Design Project Leader and all staff involved in the design and development of new products are to follow this procedure and maintain the appropriate records.

- PROCEDURE

  - Design History File

    - For each new product design or significant design change to an existing product, create a Design History File (DHF). The DHF can be hard-copy or electronic.

    - The DHF consists of the following section:

      - Design Plan

      - Design Inputs

      - Design Outputs

      - Design Review

      - Design Verification

      - Design Validation

      - Design Transfer

Note: The Design Plan, Design Inputs and Design Outputs sections are further divided into “Latest” and “Obsolete” subsections. See Section 7.9 Design Changes below for more information on control of changes as the design progresses.

  - Design Plan

    - Management assigns a Project Leader to the design project.

    - Fill out Form 7.3-1-1 Design Plan. This plan is approved by the Project Leader and/or the MR. The initial design plan starting a new design project is also approved by the President. As part of the design plan, determine if clinical investigations / trials will be performed in accordance with MDR 2017/745 Annex XV and ISO 14155, if required for Competent Authority and Ethics Committee approvals.

      - Document the following

- The design and development stages;

- The review(s) needed at each design and development stage;

- The verification, validation, and design transfer activities that are appropriate at each design and development stage;

- The responsibilities and authorizes for design and development;

- The methods to ensure traceability of design and development outputs to design and development inputs;

- The resources needed, including necessary competence of personnel.

    - Plan when Design Reviews are to take place. A Final Design Review is always done immediately before Design Transfer to production specifications. Additional Design Reviews may take place after the initial Design Inputs are complete, before and/or after completion of prototypes, after verification / validation testing, etc. At a minimum, conduct at least three Design Reviews unless otherwise justified in the Design Plan.

    - Plan the necessary Design Verification and Design Validation testing.

    - Plan Risk Management / Risk Assessment (to be updated as design progresses, see QSP 7.3-3 Risk Management).

    - Plan the Design Transfer at the completion of the design project.

    - Reference or insert any additional plans, such as Grant charts or other design planning documents, if applicable.

    - The design plans are reviewed, updated and approved as design and development evolves.

  - Design Inputs

    - Fill out Form 7.3-1-2 Design Inputs / Design Outputs with the design inputs.

- Verify that the inputs on Form 7.3-1-2 address all applicable principles (for a given design) found on the Medical Devices Essential Principles Checklist [TG(MD)R Sch3, Part 1.4(5)(III)(c)]. If an applicable principle is not already addressed, add it to the design inputs. For any principle that is not applied, provide a justification/explanation for each on Form 7.3-1-2.

Note: Principles 10, 11, 12 and 15 do not apply to current Tulip devices. Verify that these principles are not applicable to new designs.

Note: Verify that the inputs on Form 7.3-1-2 address all applicable General Safety and Performance Requirements (GSPRs) as outlined in Annex I of Regulation (EU) 2017/745.

- State of the art standards shall be used for design inputs relating to risk, biocompatibility, and sterility. The following is a list of acceptable state of the art standards:

- ISO 14971 Application of risk management to medical devices

- ISO 13485 Quality management systems: Requirements for regulatory purposes

- ISO 10993 Biological evaluation of medical devices

Note: Safety and efficacy of the device must be included in the design to include a way to verify the desired output.

    - As the design progresses, the design output column is filled out.

    - The initial Design Inputs are approved by the President and the Project Leader or MR. Subsequent changes to Form 7.3-1-2 Design Inputs / Design Outputs are approved by the Project Leader or the MR and may also be approved by the President.

      - Document the following

- Functional, performance, usability and safety requirements, according to the intended use;

- Applicable regulatory requirements and standards;

- Applicable output(s) of risk management;

- As appropriate, information derived from previous similar designs;

- Other requirements essential for design and development of the product and processes.

      - These inputs shall be reviewed for adequacy and approved.

      - Requirements shall be complete, unambiguous, able to be verified or validated, and not in conflict with each other.

Note: Further information can be found in IEC 62366-1.

      - Maintain records (see 4.2.5).

    - Place or reference Form 7.3-1-2 Design Inputs / Design Outputs and any other design input documents in the Design Inputs section of the Design History File.

  - Design Outputs

    - The finished Design Output consists of documentation that defines the device and ensures the finishing design meets the Design Input requirements. This includes, for example:

- Drawings,

- Material specifications,

- Bills of Material,

- Acceptance criteria and inspection instructions,

- Device labeling, including Instructions for Use,

- Device packaging specifications, and

- Prototypes

    - Place or reference design outputs in the Design Outputs section of the Design History File.

  - Design Reviews

    - A Design Review is a meeting where the purpose is to

      - Evaluate the ability of the results of the design to meet requirements, and

      - Identify any problems and propose necessary actions.

    - Participants at each Design Review include representative of all functions concerned with the design stage being reviewed and an individual(s) who does not have direct responsibility for the design stage being reviewed, as we as any specialists needed, if required.

    - When conducting Design Reviews, fill out Form 7.3-1-3 Design Review. Attach or include any supporting documents.

    - Place records of Design Reviews in the Design Reviews section of the Design History File.

  - Design Verification

    - The purpose of Design Verification is to confirm that Design Output meets the Design Input requirements.

    - Design Verifications results must clearly define the design (including model numbers, if applicable), the methods, the date and the individuals performing the verification.

    - The Design Verification results must also clearly state conclusions as to whether or not the Design Output meets the Design Input requirements.

    - Design Verification results must be approved by a Tulip Medical representative.

    - In the Verification plans, list the methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size of the design being evaluated.

    - All devices are required to be designed and manufactured to be safe, and particular attention paid to

      - Identifying the risks inherent in the device

      - If risks can be eliminated, eliminate them

      - If risks cannot be eliminated

        - Reduce the risks to the extent possible

        - Provide for protection appropriate to those risks,

        - Provide, with the device, information relative to the risks that remain; and

        - Minimize the hazard form the potential failures during the projected useful life of the device.

      - Verification may include, for example, fault tree analysis (FTA), failure mode and effects analysis (FMEA), product testing, performance, environmental, reliability, safety testing, software testing, electrical testing, etc.

    - If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), make sure verification includes confirmation that the design outputs meet design inputs when so connected or interfaced. Records of the results are recorded.

    - Place or reference Design Verification Results in the Design Verification section of the Design History File.

  - Design Validation

    - The purpose of Design Validation is to ensure that the device conforms to user needs and intended uses.

    - Design Validation must include testing of production units under actual or simulated use conditions. Design Validation is performed under defined operating conditions on initial production units, lots, batches, or their equivalents.

      - In the Design Validation plans, list the methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size.

      - Design Validation ensures that representative product includes initial production units, batches or their equivalents. Document the rationale for the choice of product used for validation (see 4.2.5).

      - Design and development Validation, clinical evaluations or performance evaluations of the medical device in accordance with applicable regulatory requirements must be performed and documented in accordance with standard techniques.

      - Perform Design and Development validation in accordance with regulatory requirements and document with standard techniques.

      - A device used for clinical evaluation or performance evaluation is not considered to be released for use to the customer.

      - If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), validation shall include confirmation that the requirements for the specified applicable or intended use have been met when so connected or interfaced.

      - Validation shall be completed prior to release for use of the product to the customer.

      - Records of results and conclusion of validation and necessary actions shall be maintained (see 4.2.4 and 4.2.5).

    - Design Validation results must clearly identify the design (including model numbers, id applicable), the methods, the date and the individuals performing the validation.

    - The Design Validation results must also clearly state conclusions as to whether or not user needs and intended uses were met.

    - As part of Design Validation, Tulip Medical performs clinical evaluations and/or evaluation of performance of the medical device, as required by national or regional regulations.

    - Design Validation results must be approved by a Tulip Medical representative.

    - Validation must be completed prior to the delivery or implementation of the product.

    - Clinical Investigations – Conduct clinical investigations / trials of the device in accordance with MDR 2017/745 Annex XV and ISO 14155, if required.

    - Results of validation may be recorded on Form 7.3-1-4 Design Verification / Validation Summary.

    - Place or reference Design Validation results in the Design Validation section of the Design History File.

  - Design Transfer

    - The purpose of Design Transfer is to ensure that the device design is correctly translated into production specifications. Design transfer ensures that Design Outputs are verified as suitable for manufacturing before becoming production specifications.

    - Hold and record a final Design Review immediately before Design Transfer.

    - Fill out Form 7.3-1-5 Design Transfer Checklist. The Design Transfer Checklist is approved by the President, the MR and Project Leader. The approvals signify that the design specifications are ready to be transferred to manufacturing. Approving the Design Transfer Checklist closes out the Design History File.

      - Verify Design Outputs are suitable for manufacturing before becoming final product specifications.

    - Place or reference the Design Transfer Checklist and any other Design Transfer documents in the Design Transfer section of the Design History File.

  - Design Changes

    - The Project Leader and project team members make revisions as necessary to the drawings, as well as to other Design Output documents. The Design Output documents must be signed and dated by the person who made the revision.

      - Before implementation, the changes shall be:

- Reviewed;

- Verified;

- Validated, as appropriate;

- Approved

    - Keep records of the changes to each revision of the design documents, preferably either within the document itself or referenced in the document.

    - Mark superseded design documents OBSOLETE and place or reference them in the Obsolete subsection of the appropriate section of the Design History File. Place or reference the new revision of the design document in the Latest subsection of the appropriate section of the Design history File.

    - Bring significant revisions to design documents to the attention of the Project Leader before approving. Any change to the form, fit or function of the design or which may impact the risk assessment or regulatory compliance are reviewed, verified and validated as appropriate, and approved before implementation. The review of design and development changes include evaluation of the effect of the changes on constituent parts and product already delivered if applicable. Records of such review and approval are maintained.

  - After the design project is complete:

    - Complete Form 4.2-1-2 Change Request and refer to QSP 4.2-3 for guidance. Clearly explain the reason for the change, including any related events (e.g., supplier issue, audit, test deviation) when applicable.

      - Maintain Design Files in accordance with ISO 13485 requirements - Design Planning Folder, Design Inputs/Outputs Folder, Design Review, Design Verification, Design Validation, Design Transfer, and Device Master Record

    - Create one or more Device Master Records for the product family (see QSP 7.3-2 Device Master Record).

    - Significant changes to a device design may invalidate the design criteria that the previous product was based upon. Follow this procedure from the beginning point for any major product design changes.

    - For assistance in determining if the change is significant, refer to the following guidance document links.

        - http://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/Guidance Documents/ucm080235.htm

        - http://www.meddev.info/_documents/R2_5_2-2_rev7.pdf

        - http://www.hc-sc.gc.ca/index-eng.php

- PROCESS FLOWCHART