Tulip Medical

Quality System Procedure

QSP 7.5-4

Handling, Storage, Packaging and Delivery

Approval

Vice President – Regulatory Affairs 						Date

Quality Assurance/Regulatory						Date

Production Manager							Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 05/06/2013 | Initial release. |
| 1 | 10/31/2017 | Updated for accuracy. |
| 2 | 08/29/2018 | Responsibilities added. |
| 3 | 10/01/2019 | Section 3.0 Reference Documents – added RFD-019 Performing a Drop Shipment Rev 2 Section 2.0 Scope – updated paragraph Section 4.0 Definitions – updated paragraph to reflect Single-Use products Section 7.0 Procedure – 7.2.3 added a ‘note’ explaining how product is stored to avoid a “mix-up” when stored Section 7.3 Packaging – 7.3.1 removed and/or approved supplier; added bullet points for Reusable and Single-Use Section 7.4 Delivery – 7.4.1 Corrections on who reviews and approves sales orders; 7.4.4 added ‘and other pertinent documents’; 7.4.5 added RFD-019 Performing a Drop Shipment Rev 2 ; removed 7.4.3 and added to 7.4.2 |
| 4 | 10/16/2019 | Added to Section 4.0 Definitions: Finished Goods/Product – Sterile product inventory that is fully packaged and ready to ship Added to Section 7.2 Raw Material, In-process and Finished Storage: Finished Product is stored on the mezzanine in a designated area, marked by signage and separated by a physical barrier and sign. |
| 5 | 04/01/2020 | Updated all sections for added clarity |
| 6 | 11/3/2020 | Changes to WI 1.2, WI 1.3, and WI 1.2-1, alignment of WI numbers will change in the Reference Document, procedure 7.0(7.1.2). |
| 7 | 3/12/2024 | Updated to include RFD -113 Authorized to ship list and RFD-115 EU distributer list Updated Section 7.3.2 to reference RFD-113 and RFD-115 |
| 8 | 5/23/2025 | Reviewed for MDR 2017/745 compliance. No MDD references found. Reference section updated to include MDR Annex I for alignment. Added MDR-aligned clarification to Responsibilities and Packaging sections for handling, storage, and shipping conformity under Annex I, Sections 10 & 11. |
| 9 | 07/17/2025 | Removed Management Representative from QSP 3.0 Removed nonreferenced documents Added all 3 WI’s for shipment procedures 5.0 Removed M2 and S2 areas 6.0 Changed from Vice President of Operations to Executive Vice President Updated responsibilities to combine all of the production team 7.1.2 Rephrased for clarity regarding the location and reference to product inspection compliance 7.2.2 Updated location of components 7.3 Updated subsection title for alignment with process 7.4.1 Updated correct WI number for shipping 7.4.3 Added final step in the shipping procedure for delivery |

- PURPOSE

The purpose of this document is to describe the procedures to control the handling, storing, packaging, and delivery of all product classifications whether performed at the Company or by an approved Tulip supplier.

- SCOPE

This procedure applies to Tulip Medical personnel who are directly involved with the handling, storing, packaging, and delivery of all products.

- REFERENCE DOCUMENTS

EU MDR 2017/745 – General Requirements for Storage, Packaging, and Distribution (Annex I, Chapter II, Section 10 & 11)

| QM | Quality Manual |
| --- | --- |
| QSP 6.4-1 | Cleaning Procedure Assembly Areas |
| QSP 7.2-1 | Sales Orders |
| QSP 7.4-2 | Product Inspection |
| QSP 8.2-3 | Returns and Servicing |
| QSP 8.3-1 | Control of Nonconforming Product |
| RFD-113 | Authorized to Ship List |
| RFD-115 | EU Distributor List |
| WI 7.2-1-0 | Shipping Procedure Part I of III Picking and Packing Sales Orders |
| WI 7.2-1-1 | Shipping Procedure Part II of III Shipping Labels |
| WI 7.2-1-2 | Shipping Procedure Part III of III Finalized Shipment |
| WI 5.2 | Drop Shipments |

- DEFINITIONS

Components – Product that has not been assembled into sets, kits or a reusable finished cannula.

Finished Goods/Product – Product that is fully packaged and ready to ship.

First-In/First-Out – The process by which Single-Use sterilized products are received and prioritized by expiration dates in storage.

- ABBREVIATIONS

Class I – Non-Sterile Products

Class II – Sterile Products

FIFO – First-In/First-Out

- RESPONSIBILITIES

Executive Vice President (EVP) ensures control of product distribution by approving the design and construction of packaging and shipping containers; and by documenting requirements for special conditions needed if packaging alone cannot provide preservation.

Production Team Responsible for storing, packaging, labeling, and delivering products in accordance with this procedure and related work instructions.

- PROCEDURE

  - Product Handling and Preservation – All classifications

    - Care is taken to prevent damage during handling of components and products to and from the Finished Product Storage Areas.

    - Only components and finished products that have been approved by QA in accordance with QSP 7.4-2 – Product Inspection shall be stored, and must be organized by expiration date following the First-In/First-Out (FIFO) method

    - Sterile (Class II) components and finished products expiration date is 18 months or longer when approved for distribution to OEM and Distributors.  Components and finished products with dates within the 18 months are separated and sold to non-distributors only.  Components and finished products within 6 months are removed, boxed, and noted as “obsolete”.  There are several expiration date check points:

- receiving and inspection;

- prior to assembly of sets or kits;

- when fulfilling orders;

    - Non-sterile (Class I) components has a 1-year warranty and no expiration date.

Note:	Ensure that all packaging and handling processes support compliance with MDR Annex I, Section 11.1, which requires protection against mechanical damage and contamination during transport and storage.

  - Product Storage

    - No special temperature or humidity requirements are required for product storage.

    - All component and product assembly and storage areas are to be kept clean and always organized. (QSP 6.4-1)

The sterile-packaged, single-use components and finished products storage area are separated and marked by signage and shelf barriers.

The non-sterile components stored as follows:

- Stems are stored in Reusable Component Room, in clear plastic tubes and marked accordingly

- Hubs are stored in Reusable Component Room, in covered bins and marked accordingly

- Accessories (to be engraved), are located in the Engraving area, stored in a cabinet and marked accordingly

- Hubs (to be engraved), are located in the Engraving area, stored in a cabinet and marked accordingly

The reusable, quality inspected, finished products are stored in bins in the Class I products room

    - All quality inspected, components and finished products are stored by product line, identified by product name, and identification number.

    - Product returns QSP 8.2-3 product waiting inspection, and nonconforming QSP 8.3-1 are identified and segregated until final quality inspection and/or dispositions are made.

  - Packaging For Order Fulfilment

    - Packaging operations for all finished products both single-use and reusable are performed at Tulip Medical. WI 7.2-1-0

    - Packaging product for shipping is as follows:

- Use RFD 113 Authorized to Ship List and RFD-115 EU Distributor List to determine if authorized to ship the products picked to the country you are shipping to.

- Production personnel picks/pulls product indicated on sales order and is double checked by another production employee

- Product is placed in appropriately sized shipping box with packaging material to protect the product from damage.

Note: 	Never ship cannulas in bubble mailers.

  - Delivery

    - Shipping personnel determines the best method of shipping based on the ‘priority’ indicated on the sales order. WI 7.2-1-1

    - Shipping personnel

- attaches a copy of the corresponding Shipping Label from UPS;

- and a Packing Slip from Tulip to each box that has been prepared for shipment.  For multiple boxes to the same customer, only box 1 will include the Packing List and indicated with a label “Packing List enclosed”;

- the Certificate of Compliance, SLI, and an invoice is applicable to distributors, OEMS, and DSDs, and aligns with Article 16 and Annex II of MDR to ensure proper information is provided to economic operators.

    - Production personnel shall review the shipment label and required documents for accuracy, ensuring all information is correct and consistent across all documentation prior to shipment WI 7.2-1-2

    - An approved supplier will dropship products sold by Tulip Medical directly to existing customers (WI 5.2).

- FLOWCHART

N/A