Tulip Medical

Quality System Procedure

QSP 8.4-2

Statistical Analysis

Approvals

VP Regulatory Affairs	Date

Regulatory Affairs Coordinator	Date

Document Control Administrator	Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 10/31/2018 | Initial release. |
| 1 | 5/23/2025 | Added MDR Annex I GSPR references (17.1 and 17.2) to align with MDR requirements during design verification and validation statistical assessments. Expanded Scope to include statistical methods used in receiving inspection, in-process inspection, and final inspection. Updated references to QSP 7.4-2 and added clarity regarding applicable responsibilities outside of design controls. |

- PURPOSE

To provide guidance for sample size determination and statistical analysis for testing performed in accordance with Design Controls and applicable process

- SCOPE

  - This procedure applies to all products tested for the purposes of:

    - Design verification testing and process validation testing performed as part of Design Controls,

    - Receiving inspection of incoming components and materials, and

    - In-process and final inspection of finished devices.

This procedure shall be used in conjunction with QSP 7.3-3, Risk Management, and QSP 7.4-2, Receiving and Inspection. The statistical methods outlined herein are to be used by default unless alternate methods are appropriately justified and documented. This procedure does not address statistical methods used in clinical studies.

- RESPONSIBILITIES

  - Product Development is responsible for statistical analyses for testing performed as part of design verification.

  - Operations is responsible for statistical analyses for testing performed as part of process validation.

- DEFINITIONS

  - Variables Data:  the numerical magnitude of a quality characteristic; involves reference to a continuous scale of some kind.

  - Attributes Data:  noting the presence (or absence) of some characteristic or attribute, and counting the number of units that do (or do not) possess the quality characteristic.

  - Confidence Interval: an upper and/or lower limit for how far the true mean of the entire population may vary from the sample mean (i.e., we would be CI% sure that the mean of all parts produced would fall within the upper and lower limits of the interval).

  - Reliability:  the proportion of the product required or expected to be within the specification.

  - Confidence: the probability than an interval about a sample statistic actually includes the population parameter.

  - k factor: the tolerance limit factor; the number of standard deviations that the upper and/or lower limits of the data range are away from the sample mean. A two-sided k-factor is required for data sets that are bounded by an upper and lower specification.  A one-sided k-factor is required for data sets that are bounded by either an upper specification or a lower specification.

- REFERENCES

  - Regulation (EU) 2017/745, Annex I, GSPR 17.1 and 17.2 – Requirements related to reliability and repeatability of measurement functions and analytical performance.

  - Weeks, V. Validating with Confidence:  Meeting Minimum Strength Specifications.  Medical Device & Diagnostic Industry, May 1997.

  - Gibbons, Robert D, Statistical Methods for Groundwater Monitoring, 1994, John Wiley and Sons, Inc.

  - NIST, Sematech Handbook, Section 7.2.6.3 (www.itl.nist.gov/div898/handbook/prc/section2/prc263.htm)

  - QSP xxx Design Failure Mode, Effect, and Criticality Analysis

  - Bothe, Davis R:  Measuring Process Capability, McGraw-Hill, 1997

  - Juran, J.M.; Gyrna, Frank M.:  Juran’s Quality Control Handbook, Fourth Edition, McGraw-Hill, 1988.

- PROCEDURE

  - Statistical analyses may be performed using the data analysis tool in Excel, using off-the-shelf software such as StatGraphics, or by creating a custom spreadsheet in Excel or Word.

5.2	Attributes Testing (Pass/Fail criteria)

    - Perform design and/or process FMEA in accordance with QSP 7.3-3 to identify failure modes and severity. Using the following table, assign the confidence/reliability to apply to the planned design verification or process validation testing for attributes data:

Table 5.2.1   Attributes Data

| Severity (S) of Harm | Severity | Confidence | Reliability |
| --- | --- | --- | --- |
| Rank/Category |  |  |  |
| 1 - Insignificant | Little or no potential for injury. Cosmetic | 90% | 90% |
| 2 - Slight | Little or no potential for injury; minor change in performance | 95% | 90% |
| 3 - Moderate | Potential for injury; major dissatisfaction with the device performance | 95% | 90% |
| 4 - Critical | Potential for serious injury or surgical intervention | 95% | 90% |
| 5 - Catastrophic | Potential death or serious injury that is life threatening | 95% | 95% |

    - When analyzing attributes data, the exact binomial distribution may be used to estimate the proportion of successes in the population.

    - The following is an example of the verbiage that may be used to express the results of the analysis of attributes data:

The statistical analysis applied to this test is based on a simple passing or failing the test criteria.  Using the exact binomial distribution with 30 units and 0 failures, with 95% confidence at least 90.5% of the population would meet the inspection criteria.

    - As a guide, the table in Appendix 1 provides the proportion of successes in the population with 95% confidence for various sample sizes, with 0, 1, 2, and 3 failures.

  - Variables Data (Continuous Data)

    - Perform design and/or process FMEA in accordance with QSP 7.3-3 to identify failure modes and severity. Using Table 5.3.1, assign the confidence/reliability to apply to the planned design verification or process validation testing for attributes data:

    - When analyzing variables data, typically two methods (described in sections 5.3.4 and 5.3.5) are used to determine if the data meet the acceptance criteria with statistical significance.  Both methods are valid provided a normal distribution can be assumed.

    - If the data are not normally distributed, either:

      - Transform the data by replacing the variable with its logarithm, inverse, or square root; then determine if the transformed data are normal.

      - Convert the variables data to attributes data. Converting to attributes may require additional samples to meet the confidence/reliability requirements. Confidence/reliability requirements may change by converting to attributes data. See table 5.2.1 above.

        - If testing requirements change from variables to attributes, the reasoning must be explained in the test report.

Table 5.3.1    Variables Data

| Severity (S) of Harm | Severity | Confidence | Reliability |
| --- | --- | --- | --- |
| Rank/Category |  |  |  |
| 1 - Insignificant | Little or no potential for injury. Cosmetic | 95% | 90% |
| 2 - Slight | Little or no potential for injury; minor change in performance | 95% | 90% |
| 3 - Moderate | Potential for injury; major dissatisfaction with the device performance | 95% | 95% |
| 4 - Critical | Potential for serious injury or surgical intervention | 95% | 95% |
| 5 - Catastrophic | Potential death or serious injury that is life threatening | 95% | 99% |

5.3.3	Process Capability is the measured, inherent reproducibility of the product turned out by a process. The Performance Index, Cpk, reflects the proximity of the current process mean to either the upper specification limit or lower specification limit.  The Cpk may be calculated to determine if a process is capable of consistently meeting the acceptance criteria.

      - Two-Sided Specification (ex:  5.0mm ± 0.25mm)

5.3.4.2	One-Sided Lower Spec Limit (ex: min 0.44N)

      - One-Sided Upper Spec Limit (ex: max 0.26”)

where  is the sample mean, s is the sample standard deviation

LSL is the lower spec limit, and USL is the upper spec limit.

The calculated Cpk must be equal to or greater than 1.0 for acceptance.

    - Statistical Tolerance Limits are similar to process capability in that they show the practical boundaries of process variability. While process capability is based on 6 practices, tolerance limits allow for decisions to be made regarding confidence levels and population percentages.

The equation for calculating statistical tolerance limits is , where k is the tolerance limit factor.

The table in Appendix 2 provides factors k such that the probability (confidence) is  that at least a proportion P of the distribution (reliability) will be less than  or greater than , where and s are estimates of the mean and standard deviation computed from a sample size of n.

The table in Appendix 3 provides two-sided factors k which is applicable to those data sets bounded by an upper and lower specification: .  See your Quality Engineer team member if the table does not provide the k-factor needed for your analysis.

  - Hypothesis Testing

    - Table 5.4.1 provides guidance for proper use of Tests of Hypotheses.

Table 5.4.1   Tests of Hypotheses

| Test of Hypothesis | Comparison | Reference Distribution |
| --- | --- | --- |
| t test of means | Compare the AVERAGE of a TEST GROUP to a STANDARD VALUE | Standard Normal Distribution when  is known Student’s t Distribution when  is estimated by s |
|  | Compare the AVERAGE of ONE TEST GROUP to the AVERAGE of a SECOND TEST GROUP |  |
|  | Compare the FIRST HALF of PAIRED DATA to the SECOND HALF |  |
| 2 test of variance | Compare the VARIANCE of a TEST GROUP to a STANDARD VALUE | 2 Distribution |
| F test of variance | Compare the VARIANCE of ONE TEST GROUP to the VARIANCE of a SECOND TEST GROUP | F Distribution |

    - Procedure for Hypothesis Testing

      - The null and alternate hypotheses are stated to clarify the purpose of the test

      - The assumptions required for validity of the statistical method are stated

Example:  Assume the data are normal and independent or Assume the variances are equivalent.

      - The confidence level of the test is specified (typically  = 0.05)

      - The test statistic is calculated

      - The critical value of the test statistic is found from the appropriate reference distribution

      - The calculated value of the test statistic is compared to the critical value as a criterion for acceptance or rejection of the alternate hypothesis.

    - At ev3 Neurovascular, the most common comparison will be between the average of one test group and the average of a second test group with  estimated by s (Student’s t test). The formulas to follow may be found below.

H0:  1 = 2

HA:  1  2  or 1>2

or      since  is estimated by s

Where  is the population average, is the average of the data set,  is the standard deviation of the population, s is the standard deviation of the data set, and n is the number of samples in the data set.

| HA | Critical Value (cv)  known | Accept HA if: |
| --- | --- | --- |
| 12 | ± t/2 | tcalc < -cv or tcalc > cv |
| 1>2 | + t | tcalc > cv |

If the alternate hypothesis is accepted, the null hypothesis is rejected.  Conversely, if the alternate hypothesis is rejected, the null hypothesis is accepted.

5.4.4	The table in Appendix 4 provides some values for the t distribution.

- APPENDICES

Appendix 1	Binomial Distribution

Appendix 2	One Sided k-factors (from Juran’s Handbook, Fourth Edition, Table AII.36)

Appendix 3	Two Sided k-factors (from Juran’s Handbook, Fourth Edition, Table AII.37)

Appendix 4	t Distribution (from Juran’s Handbook, Fourth Edition, Table AII.16)

Appendix 5	2 Distribution (from Measuring Process Capability, Table VI)

Appendix 1			 Binomial Distribution with 95% Confidence

| Sample Size | 0 Failures | 1 Failure | 2 Failures | 3 Failures |
| --- | --- | --- | --- | --- |
| 15 | 81.9 | 72.1 | 63.7 | 56.0 |
| 20 | 86.1 | 78.4 | 71.7 | 65.6 |
| 25 | 88.7 | 82.4 | 76.9 | 71.8 |
| 26 | 89.1 | 83.0 | 77.7 | 72.8 |
| 27 | 89.5 | 83.6 | 78.5 | 73.7 |
| 28 | 89.9 | 84.1 | 79.2 | 74.6 |
| 29 | 90.2 | 84.7 | 79.8 | 75.4 |
| 30 | 90.5 | 85.1 | 80.5 | 76.1 |
| 35 | 91.8 | 87.1 | 83.1 | 79.3 |
| 40 | 92.8 | 88.7 | 85.1 | 81.7 |
| 45 | 93.6 | 89.9 | 86.7 | 83.7 |
| 50 | 94.2 | 90.9 | 87.9 | 85.2 |
| 55 | 94.7 | 91.7 | 89.0 | 86.5 |
| 56 | 94.8 | 91.8 | 89.2 | 86.7 |
| 57 | 94.9 | 91.9 | 89.4 | 87.0 |
| 58 | 95.0 | 92.1 | 89.5 | 87.2 |
| 59 | 95.0 | 92.2 | 89.7 | 87.4 |
| 60 | 95.1 | 92.3 | 89.9 | 87.6 |
| 65 | 95.5 | 92.9 | 90.6 | 88.5 |
| 70 | 95.8 | 93.4 | 91.3 | 89.3 |
| 75 | 96.1 | 93.8 | 91.8 | 90.0 |
| 100 | 97.0 | 95.3 | 93.8 | 92.4 |

*See the Quality Engineering team member if additional information is required regarding the binomial distribution.

Appendix 2				Statistical Tolerance Limits

| One-Sided Factors |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| n P |  = 0.90 |  |  |  |  = 0.95 |  |  |  |
|  | 0.90 | 0.95 | 0.99 | 0.999 | 0.90 | 0.95 | 0.99 | 0.999 |
| 3 | 4.258 | 5.310 | 7.340 | 9.651 | 6.158 | 7.655 | 10.552 | 13.857 |
| 4 | 3.187 | 3.957 | 5.437 | 7.128 | 4.163 | 5.145 | 7.042 | 9.215 |
| 5 | 2.742 | 3.400 | 4.666 | 6.112 | 3.407 | 4.202 | 5.741 | 7.501 |
| 6 | 2.494 | 3.091 | 4.242 | 5.556 | 3.006 | 3.707 | 5.062 | 6.612 |
| 7 | 2.333 | 2.894 | 3.972 | 5.201 | 2.755 | 3.399 | 4.641 | 6.061 |
| 8 | 2.219 | 2.755 | 3.783 | 4.955 | 2.582 | 3.188 | 4.353 | 5.686 |
| 9 | 2.133 | 2.649 | 3.641 | 4.772 | 2.454 | 3.031 | 4.143 | 5.414 |
| 10 | 2.065 | 2.568 | 3.532 | 4.629 | 2.355 | 2.911 | 3.981 | 5.203 |
| 11 | 2.012 | 2.503 | 3.444 | 4.515 | 2.275 | 2.815 | 3.852 | 5.036 |
| 12 | 1.966 | 2.448 | 3.371 | 4.420 | 2.210 | 2.736 | 3.747 | 4.900 |
| 13 | 1.928 | 2.403 | 3.310 | 4.341 | 2.155 | 2.670 | 3.659 | 4.787 |
| 14 | 1.895 | 2.363 | 3.257 | 4.274 | 2.108 | 2.614 | 3.585 | 4.690 |
| 15 | 1.866 | 2.329 | 3.212 | 4.215 | 2.068 | 2.566 | 3.520 | 4.607 |
| 16 | 1.842 | 2.299 | 3.172 | 4.164 | 2.032 | 2.523 | 3.463 | 4.534 |
| 17 | 1.820 | 2.272 | 3.136 | 4.118 | 2.001 | 2.486 | 3.415 | 4.471 |
| 18 | 1.800 | 2.249 | 3.106 | 4.078 | 1.974 | 2.453 | 3.370 | 4.415 |
| 19 | 1.781 | 2.228 | 3.078 | 4.041 | 1.949 | 2.423 | 3.331 | 4.364 |
| 20 | 1.765 | 2.208 | 3.052 | 4.009 | 1.926 | 2.396 | 3.295 | 4.319 |
| 21 | 1.750 | 2.190 | 3.028 | 3.979 | 1.905 | 2.371 | 3.262 | 4.276 |
| 22 | 1.736 | 2.174 | 3.007 | 3.952 | 1.887 | 2.350 | 3.233 | 4.238 |
| 23 | 1.724 | 2.159 | 2.987 | 3.927 | 1.869 | 2.329 | 3.206 | 4.204 |
| 24 | 1.712 | 2.145 | 2.969 | 3.904 | 1.853 | 2.309 | 3.181 | 4.171 |
| 25 | 1.702 | 2.132 | 2.952 | 3.882 | 1.838 | 2.292 | 3.158 | 4.143 |
| 30 | 1.657 | 2.080 | 2.884 | 3.794 | 1.778 | 2.220 | 3.064 | 4.022 |
| 35 | 1.623 | 2.041 | 2.833 | 3.730 | 1.732 | 2.166 | 2.994 | 3.934 |
| 40 | 1.598 | 2.010 | 2.793 | 3.679 | 1.697 | 2.126 | 2.941 | 3.866 |
| 45 | 1.577 | 1.986 | 2.762 | 3.638 | 1.669 | 2.092 | 2.897 | 3.811 |
| 50 | 1.560 | 1.965 | 2.735 | 3.604 | 1.646 | 2.065 | 2.863 | 3.766 |

Appendix 3				Statistical Tolerance Limits

| Two-Sided Factors |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| n P |  = 0.90 |  |  |  |  = 0.95 |  |  |  |
|  | 0.90 | 0.95 | 0.99 | 0.999 | 0.90 | 0.95 | 0.99 | 0.999 |
| 2 | 15.978 | 18.800 | 24.167 | 30.227 | 32.019 | 37.674 | 48.430 | 60.573 |
| 3 | 5.847 | 6.919 | 8.974 | 11.309 | 8.380 | 9.916 | 12.861 | 16.208 |
| 4 | 4.116 | 4.943 | 6.440 | 8.149 | 5.369 | 6.370 | 8.299 | 10.502 |
| 5 | 3.494 | 4.152 | 5.423 | 6.879 | 4.275 | 5.079 | 6.634 | 8.415 |
| 6 | 3.131 | 3.723 | 4.870 | 6.188 | 3.712 | 4.414 | 5.775 | 7.337 |
| 7 | 2.902 | 3.452 | 4.521 | 5.750 | 3.369 | 4.007 | 5.248 | 6.676 |
| 8 | 2.743 | 3.264 | 4.278 | 5.446 | 3.136 | 3.732 | 4.891 | 6.226 |
| 9 | 2.626 | 3.125 | 4.098 | 5.220 | 2.967 | 3.532 | 4.631 | 5.899 |
| 10 | 2.535 | 3.018 | 3.959 | 5.046 | 2.839 | 3.379 | 4.433 | 5.649 |
| 11 | 2.463 | 2.933 | 3.849 | 4.906 | 2.737 | 3.259 | 4.277 | 5.452 |
| 12 | 2.404 | 2.863 | 3.758 | 4.792 | 2.655 | 3.162 | 4.150 | 5.291 |
| 13 | 2.355 | 2.805 | 3.682 | 4.697 | 2.587 | 3.081 | 4.044 | 5.158 |
| 14 | 2.314 | 2.756 | 3.618 | 4.615 | 2.529 | 3.012 | 3.955 | 5.045 |
| 15 | 2.278 | 2.713 | 3.562 | 4.545 | 2.480 | 2.954 | 3.878 | 4.949 |
| 16 | 2.246 | 2.676 | 3.514 | 4.484 | 2.437 | 2.903 | 3.812 | 4.865 |
| 17 | 2.219 | 2.643 | 3.471 | 4.430 | 2.400 | 2.858 | 3.754 | 4.791 |
| 18 | 2.194 | 2.614 | 3.433 | 4.382 | 2.366 | 2.819 | 3.702 | 4.725 |
| 19 | 2.172 | 2.588 | 3.399 | 4.339 | 2.337 | 2.784 | 3.656 | 4.667 |
| 20 | 2.152 | 2.564 | 3.368 | 4.300 | 2.310 | 2.752 | 3.615 | 4.614 |
| 21 | 2.135 | 2.543 | 3.340 | 4.264 | 2.286 | 2.723 | 3.577 | 4.567 |
| 22 | 2.118 | 2.524 | 3.315 | 4.232 | 2.264 | 2.697 | 3.543 | 4.523 |
| 23 | 2.103 | 2.506 | 3.292 | 4.203 | 2.244 | 2.673 | 3.512 | 4.484 |
| 24 | 2.089 | 2.480 | 3.270 | 4.176 | 2.225 | 2.651 | 3.483 | 4.447 |
| 25 | 2.077 | 2.474 | 3.251 | 4.151 | 2.208 | 2.631 | 3.457 | 4.413 |
| 26 | 2.065 | 2.460 | 3.232 | 4.127 | 2.193 | 2.612 | 3.432 | 4.382 |
| 27 | 2.054 | 2.447 | 3.215 | 4.106 | 2.178 | 2.595 | 3.409 | 4.353 |
| 30 | 2.025 | 2.413 | 3.170 | 4.049 | 2.140 | 2.549 | 3.350 | 4.278 |
| 35 | 1.988 | 2.368 | 3.112 | 3.974 | 2.090 | 2.490 | 3.272 | 4.179 |
| 40 | 1.959 | 2.334 | 3.066 | 3.917 | 2.052 | 2.445 | 3.213 | 4.104 |
| 45 | 1.935 | 2.306 | 3.030 | 3.871 | 2.021 | 2.408 | 3.165 | 4.042 |
| 50 | 1.916 | 2.284 | 3.001 | 3.833 | 1.996 | 2.379 | 3.126 | 3.993 |

Appendix 4						t Distribution

| n | t.90 | t.95 | t.975 | t.99 | t.995 |
| --- | --- | --- | --- | --- | --- |
| 1 | 3.078 | 6.314 | 12.706 | 31.821 | 63.657 |
| 2 | 1.886 | 2.920 | 4.303 | 6.965 | 9.925 |
| 3 | 1.638 | 2.353 | 3.182 | 4.541 | 5.841 |
| 4 | 1.533 | 2.132 | 2.776 | 3.747 | 4.604 |
| 5 | 1.476 | 2.015 | 2.571 | 3.365 | 4.032 |
| 6 | 1.440 | 1.943 | 2.447 | 3.143 | 3.707 |
| 7 | 1.415 | 1.895 | 2.365 | 2.998 | 3.499 |
| 8 | 1.397 | 1.860 | 2.306 | 2.896 | 3.355 |
| 9 | 1.383 | 1.833 | 2.262 | 2.821 | 3.250 |
| 10 | 1.372 | 1.812 | 2.228 | 2.764 | 3.169 |
| 11 | 1.363 | 1.796 | 2.201 | 2.718 | 3.106 |
| 12 | 1.356 | 1.782 | 2.179 | 2.681 | 3.055 |
| 13 | 1.350 | 1.771 | 2.160 | 2.650 | 3.012 |
| 14 | 1.345 | 1.761 | 2.145 | 2.624 | 2.977 |
| 15 | 1.341 | 1.753 | 2.131 | 2.602 | 2.947 |
| 16 | 1.337 | 1.746 | 2.120 | 2.583 | 2.921 |
| 17 | 1.333 | 1.740 | 2.110 | 2.567 | 2.898 |
| 18 | 1.330 | 1.734 | 2.101 | 2.552 | 2.878 |
| 19 | 1.328 | 1.729 | 2.093 | 2.539 | 2.861 |
| 20 | 1.325 | 1.725 | 2.086 | 2.528 | 2.845 |
| 21 | 1.323 | 1.721 | 2.080 | 2.518 | 2.831 |
| 22 | 1.321 | 1.717 | 2.074 | 2.508 | 2.819 |
| 23 | 1.319 | 1.714 | 2.069 | 2.500 | 2.807 |
| 24 | 1.318 | 1.711 | 2.064 | 2.492 | 2.797 |
| 25 | 1.316 | 1.708 | 2.060 | 2.485 | 2.787 |
| 26 | 1.315 | 1.706 | 2.056 | 2.479 | 2.779 |
| 27 | 1.314 | 1.703 | 2.052 | 2.473 | 2.771 |
| 28 | 1.313 | 1.701 | 2.048 | 2.467 | 2.763 |
| 29 | 1.311 | 1.699 | 2.045 | 2.462 | 2.756 |
| 30 | 1.310 | 1.697 | 2.042 | 2.457 | 2.750 |
| 40 | 1.303 | 1.684 | 2.021 | 2.423 | 2.704 |
| 60 | 1.296 | 1.671 | 2.000 | 2.390 | 2.660 |
| 120 | 1.289 | 1.658 | 1.980 | 2.358 | 2.617 |
|  | 1.282 | 1.645 | 1.960 | 2.326 | 2.576 |

 = degrees of freedom = min(n1,n2)-1

Appendix 5					Chi-Squared Distribution

|  | =.10 | =.05 | =.01 |  |  | =.10 | =.05 | =.01 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 2.706 | 3.841 | 6.635 |  | 51 | 64.295 | 68.669 | 77.386 |
| 2 | 4.605 | 5.991 | 9.210 |  | 52 | 65.422 | 69.832 | 78.616 |
| 3 | 6.251 | 7.815 | 11.345 |  | 53 | 66.548 | 70.993 | 79.843 |
| 4 | 7.779 | 9.488 | 13.277 |  | 54 | 67.673 | 72.153 | 81.069 |
| 5 | 9.236 | 11.071 | 15.086 |  | 55 | 68.796 | 73.311 | 82.292 |
| 6 | 10.645 | 12.592 | 16.812 |  | 56 | 69.919 | 74.468 | 83.513 |
| 7 | 12.017 | 14.067 | 18.475 |  | 57 | 71.040 | 75.624 | 84.733 |
| 8 | 13.362 | 15.507 | 20.090 |  | 58 | 72.160 | 76.778 | 85.950 |
| 9 | 14.684 | 16.919 | 21.666 |  | 59 | 73.279 | 77.931 | 87.166 |
| 10 | 15.987 | 18.307 | 23.209 |  | 60 | 74.397 | 79.082 | 88.379 |
| 11 | 17.275 | 19.675 | 24.725 |  | 61 | 75.514 | 80.232 | 89.591 |
| 12 | 18.549 | 21.026 | 26.217 |  | 62 | 76.630 | 81.381 | 90.802 |
| 13 | 19.812 | 22.362 | 27.688 |  | 63 | 77.745 | 82.529 | 92.010 |
| 14 | 21.064 | 23.685 | 29.141 |  | 64 | 78.860 | 83.675 | 93.217 |
| 15 | 22.307 | 24.996 | 30.578 |  | 65 | 79.973 | 84.821 | 94.422 |
| 16 | 23.542 | 26.296 | 32.000 |  | 66 | 81.085 | 85.965 | 95.626 |
| 17 | 24.769 | 27.587 | 33.409 |  | 67 | 82.197 | 87.108 | 96.828 |
| 18 | 25.989 | 28.869 | 34.805 |  | 68 | 83.308 | 88.250 | 98.028 |
| 19 | 27.204 | 30.144 | 36.191 |  | 69 | 84.418 | 89.391 | 99.228 |
| 20 | 28.412 | 31.410 | 37.566 |  | 70 | 85.527 | 90.531 | 100.425 |
| 21 | 29.615 | 32.671 | 38.932 |  | 71 | 86.635 | 91.670 | 101.621 |
| 22 | 30.813 | 33.924 | 40.289 |  | 72 | 87.743 | 92.808 | 102.816 |
| 23 | 32.007 | 35.172 | 41.638 |  | 73 | 88.850 | 93.945 | 104.010 |
| 24 | 33.196 | 36.415 | 42.980 |  | 74 | 89.956 | 95.081 | 105.202 |
| 25 | 34.382 | 37.652 | 44.314 |  | 75 | 91.061 | 96.217 | 106.393 |
| 26 | 35.563 | 38.885 | 45.642 |  | 76 | 92.166 | 97.351 | 107.583 |
| 27 | 36.741 | 40.113 | 46.963 |  | 77 | 93.270 | 98.484 | 108.771 |
| 28 | 37.916 | 41.337 | 48.278 |  | 78 | 94.374 | 99.617 | 109.958 |
| 29 | 39.087 | 42.557 | 49.588 |  | 79 | 95.476 | 100.749 | 111.144 |
| 30 | 40.256 | 43.773 | 50.892 |  | 80 | 96.578 | 101.879 | 112.329 |
| 31 | 41.422 | 44.985 | 52.191 |  | 81 | 97.680 | 103.010 | 113.512 |
| 32 | 42.585 | 46.194 | 53.486 |  | 82 | 98.780 | 104.139 | 114.695 |
| 33 | 43.745 | 47.400 | 54.776 |  | 83 | 99.880 | 105.267 | 115.876 |
| 34 | 44.903 | 48.602 | 56.061 |  | 84 | 100.980 | 106.395 | 117.057 |
| 35 | 46.059 | 49.802 | 57.342 |  | 85 | 102.079 | 107.522 | 118.236 |
| 36 | 47.212 | 50.998 | 58.619 |  | 86 | 103.177 | 108.648 | 119.414 |
| 37 | 48.363 | 52.192 | 59.893 |  | 87 | 104.275 | 109.773 | 120.591 |
| 38 | 49.513 | 53.384 | 61.162 |  | 88 | 105.372 | 110.898 | 121.767 |
| 39 | 50.660 | 54.572 | 62.428 |  | 89 | 106.469 | 112.022 | 122.942 |
| 40 | 51.805 | 55.758 | 63.691 |  | 90 | 107.565 | 113.145 | 124.116 |
| 41 | 52.949 | 56.942 | 64.950 |  | 91 | 108.661 | 114.268 | 125.289 |
| 42 | 54.090 | 58.124 | 66.206 |  | 92 | 109.756 | 115.390 | 126.462 |
| 43 | 55.230 | 59.304 | 67.459 |  | 93 | 110.850 | 116.511 | 127.633 |
| 44 | 56.369 | 60.481 | 68.710 |  | 94 | 111.944 | 117.632 | 128.803 |
| 45 | 57.505 | 61.656 | 69.957 |  | 95 | 113.038 | 118.752 | 129.973 |
| 46 | 58.641 | 62.830 | 71.201 |  | 96 | 114.131 | 119.871 | 131.141 |
| 47 | 59.774 | 64.001 | 72.443 |  | 97 | 115.223 | 120.990 | 132.309 |
| 48 | 60.907 | 65.171 | 73.683 |  | 98 | 116.315 | 122.108 | 133.476 |
| 49 | 62.038 | 66.339 | 74.919 |  | 99 | 117.407 | 123.225 | 134.642 |
| 50 | 63.167 | 67.505 | 76.154 |  | 100 | 118.498 | 124.342 | 135.807 |

 = degrees of freedom, typically 2