Tulip Medical

Quality System Procedure

QSP 8.4-3

Statistical Sampling Plan for Verification & Validation

Approvals

V.P Regulatory Affairs 						             Date

Regulatory Affairs Coordinator					             Date

Document Control Administrator					              Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 10/14/2019 | Initial Release |
| 1 | 10/19/2020 | Changed “Reliability” and “Sample Size” for level 1 severity situations. They are now specified on a case-by-case basis by R&D. Changed font to Calibri |
| 2 | 5/23/2025 | Added reference to Regulation (EU) 2017/745 on medical devices (Annex I, II, III) in the REFERENCES section to support alignment with MDR requirements for design and process validation activities. Optional MDCG guidance (2020-10/1) was also included to enhance traceability to clinical evaluation practices. |

- PURPOSE

To provide sampling plans for verification & validation testing along with the statistical analysis behind the sampling plans

- SCOPE

This procedure applies to all products tested for the purpose of design & process verification/validation. The sampling plans in this procedure shall be used by default unless more appropriate methods are justified. This procedure does not provide sampling plans for clinical studies or ongoing quality control testing.

- RESPONSIBILITIES

Product Development is responsible for using sampling plans for verification/validation testing performed as part of design & process

- REFERENCES

Regulation (EU) 2017/745 on medical devices – Annex I (GSPR), Annex II & III

QSP 8.4-1	Testing, Measurement and Analysis

RSK-04 		Process FMEA

RSK-09 		Design FMEA

Montgomery DC. Introduction to Statistical Quality Control. Sixth Ed. Hoboken, New Jersey:

John Wiley & Sons, Inc; 2009

MDCG 2020-10/1 – Clinical Evaluation: Guide for manufacturers and notified bodies

- DEFINITIONS

Variable Data – the numerical magnitude of a quality characteristic; involves reference to a continuous scale of some kind.

Attribute Data – the sum of units that do or do not possess the presence or absence of some characteristic or attribute.

Confidence – the probability that an interval about a sample statistic includes the population parameter.

Reliability – the proportion of the product expected to be within specification.

- PROCEDURE

- Identify the type of data that will be collected during testing (i.e. attribute or variable).

- Identify the level of severity associated with the test sample using RSK-09.

- Use Table 6.3 to determine the sample size to be used in testing based upon severity.

| Table 6.3 – Sampling Plan for Attribute & Variable Data |  |  |  |  |
| --- | --- | --- | --- | --- |
| Severity of Harm | Confidence | Reliability | Sample Size | Max Number of Failures |
| 1 - Insignificant | 95% | Specified by R&D | Specified by R&D | 0 |
|  |  |  | Specified by R&D | 1 |
|  |  |  | Specified by R&D | 2 |
|  |  |  | Specified by R&D | 3 |
| 2 - Slight | 95% | 85% | 19 | 0 |
|  |  |  | 30 | 1 |
|  |  |  | 40 | 2 |
|  |  |  | 50 | 3 |
| 3 - Moderate | 95% | 90% | 29 | 0 |
|  |  |  | 46 | 1 |
|  |  |  | 61 | 2 |
|  |  |  | 76 | 3 |
| 4 - Critical | 95% | 95% | 59 | 0 |
|  |  |  | 93 | 1 |
|  |  |  | 124 | 2 |
|  |  |  | 153 | 3 |
| 5 - Catastrophic | 95% | 99% | 299 | 0 |
|  |  |  | 473 | 1 |
|  |  |  | 628 | 2 |
|  |  |  | 773 | 3 |

- SAMPLE SIZE EQUATION

- Equation used to determine sample sizes for attribute & variable data (based on binomial distribution w/ infinite population):

β≥i=0Cn!i!n-i!1-RLiRLn-i

where   n = sample size

β = 1 – confidence

RL = Required Reliability

C = number of allowable failures