Tulip Medical

Quality System Procedure

QSP 9.1-6

Medical Device Reporting

Approvals

10/12/2018

Director of Operations 						Date

10/12/2018

Management Representative 						Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 05/06/2013 | Release documentation – Implementation of document for compliance to FDA 21 CFR 803 Medical Device Reporting |
| 1 | 09/18/2014 | Updated references to ISO 13485:2012 |
| 2 | 07/16/2015 | Updated Paragraph 6.4 for reporting requirements timeline to FDA, Canada and EU. |
| 3 | 09/11/2018 | Reformatted and simplified procedure. New definitions added under Section 4.0. Addition of Section 8.0 Process Flowchart. Replacement of ‘COO’ with ‘Operations Management.’ |

- PURPOSE

To provide methods and forms used for reporting under the FDA’s Medical Device Reporting regulation Title 21 CFR, Part 803. Encompasses the FDA reporting and record-keeping requirements for information received, which reasonable suggests that a manufactured device may have caused, or contributed to, a death, serious injury, or has malfunctioned, or that the device would likely cause, or contribute to, a death or serious injury if the malfunction were to recur.

- SCOPE

Although administrative requirements may vary by global region, this procedure applies to Tulip products sold anywhere in the world.

- REFERENCE DOCUMENTS

| Food and Drug Administration 21 CFR 820 Quality System Regulation 820.100 |  |
| --- | --- |
| Food and Drug Administration 21 CFR 803 Medical Device Reporting |  |
| International Standard Medical devices – Quality management systems – Requirements for regulatory purposes ISO 13485:2012 |  |
| Form 6.2-1-2 | Training Record |
| Form 8.2-2-2-2 | Complaint Investigation |
| QSP 6.2-1 | Training and Qualification |
| QSP 8.2-2 | Customer Feedback / Complaints |

- DEFINITIONS

Becomes Aware – The moment an employee of the company has acquired information reasonably suggesting that a reportable adverse event has occurred.

Caused/Contributed – A death or serious injury was, or may have been, attributed to a medical device, or that may have been a factor in a death or serious injury. This includes events occurring as a result of device failure, malfunction, improper design, manufacturing defect, labeling error, or user error.

Complaint – Any written, electronic, or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety, effectiveness or performance of a device after it is released for distribution.

Malfunction – The failure of the device to meet its performance specifications, or otherwise perform as intended. A malfunction should be considered reportable if any of the following is true.

- The chance of it causing or contributing to a death or serious injury is not remote or minute.

- It affects the device in a catastrophic manner that may lead to a death or serious injury.

- It causes the device to fail to perform its essential functions and compromises the device’s therapeutic, monitoring or diagnostic effectiveness, which could cause or contribute to a death or serious injury.

- The device involves a long-term implant or device that is considered to be life supporting, or life sustaining, and thus essential to maintaining human life.

- The manufacturer takes, or would be required t take, action to reduce a risk to health as a result of the malfunction.

- A malfunction of the same type has actually caused or contributed to a death or serious injury in the past.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (QSPs).

MDR Reportable Event – An event about which the company has received or becomes aware of information that reasonably suggests that one of our marketed devices:

- May have caused or contributed to a death or serious injury; or

- Has malfunctioned and that the device, or a similar device marketed, would be likely to cause or contribute to a death or serious injury if the malfunction were to recur.

Reasonably Suggests – FDA defines this as, “A report is required when a reporting entity receives information that ‘reasonably suggests’ that a device may have caused or contributed to a reportable event. Information that ‘reasonably suggests’ includes any information such as professional, scientific or medical facts and observations or opinions that would reasonable suggest that a device may have contributed to a reportable event.”

Remedial Action – Any action, other than routine maintenance or servicing of a device, where such action is necessary to prevent recurrence of a reportable event.

Serious Injury – An injury or illness that is:

  - Life-threatening,

  - Results in permanent impairment to a body function or permanent damage to body structure, or

  - Necessitates medical or surgical intervention to preclude permanent impairment of a body function or permanent damage to a body structure.

- ABBREVIATIONS

FDA – Food and Drug Administration

MDR – Medical Device Report

MR – Management Representative

OM – Operations Management

- RESPONSIBILITIES

The Management Representative (MR) is responsible for maintaining this procedure in a state current with applicable regulations and for ensuring that the company and its representatives follow their respective reporting responsibilities. They are responsible for ensuring that personnel qualified to carry out adverse event reporting responsibilities within their respective regions represent Tulip products outside of the US.

Authorized Representatives outside of the US are responsible for adverse event reporting in their respective regions.

- PROCEDURE

  - Report Determination

    - If it is determined that an event is to be reported, all the regions affected shall be identified. If reporting is required in more than one region, the Management Representative (MR) shall coordinate all the reporting activities across the regions.

    - If an investigation determines that an event does not need to be reported, the decision shall be documented on Form 8.2-2-2-2 Complaint Investigation.

    - Reports must be made in a timely manner. As reporting timeframes may vary by region, it is the responsibility of the regional representative to comply with country specific requirements.

  - Submitting Reports

    - Each report shall be submitted in accordance with the requirements of the regions in which they are being reported.

    - All reports outside the US shall be done with the approval of the MR.

  - Quality Records

    - All correspondence related to a MDR submission shall be kept with the applicable Complaint file and clearly identified as a MDR, or segregated as a MDR file.

    - The MDR files must contain an explanation of why any information required by the MDR regulation was not submitted, or could not be obtained.

    - The complete results of any device investigation or evaluation are to be documented and maintained as part of the MDR adverse event file.

    - The MDR files must also include information on any discussions and the decision-making processes used to determine if a device related death, serious injury, or malfunction was or was not reportable.

- PROCESS FLOWCHART