Tulip Medical

Quality System Procedure

QSP 9.1-7

Post-Market Surveillance and Post-Market Clinical Follow-up

Approvals

V.P Regulatory Affairs 						             	Date

Regulatory Affairs Coordinator					             	Date

Document Control Administrator					Date

| Rev | Effective Date | Description of Change | Originator |
| --- | --- | --- | --- |
| 0 | 05/06/2013 | Released documentation. |  |
| 1 | 05/15/2014 | Updated references for ISO 13485:2012. |  |
| 2 | 11/15/2014 | Updated duties of the Sales Team. |  |
| 3 | 09/11/2018 | Reformatted document and replaced ‘COO’ with ‘EVP.’ |  |
| 4 | 05/01/2019 | Section 3.0 added references; Section 4.0 added Definitions; Section 7.1 Inserted FDA Website for MUU Classification; Section 7.6 Reporting Results; Section 7.6.2 Requirements of Article 87 of the MDR and MEDDEV 2.12-1 Rev 8. Defined Post Market Surveillance Plan and Report requirements. |  |
| 5 | 7/9/2019 | All references to RFD-068, RFD-069, RFD-070, and RFD-071 will now refer to PM-001-C, PM-001-R, PM-001-Reuse, and PM-002-Single |  |
| 6 | 7/31/2019 | Sections 7.6.1 Removed all references to PSUR, Periodic Safety Update Report. References to a PMSR were removed, Post Market Surveillance Report for both Class I and Class II devices will substitute for the PSUR. |  |
| 7 | 5/1/2023 | Release documentation: - Implementation of document for EU Medical Device Regulations (MDR) 2017/745 for Class Ila Devices - PMS requirements from Chapter VII, Sec. 1, Art. 83, 84, 85 & Annex III & Annex XIV Part B. Release of forms: Form 9.1-3-1, 9.1-3-2, 9.1-3-3, 9.1-3-4, 9.1-3-5. Added References US FDA 21 CFR Part 806 , US FDA 21 CFR Part 7, Health Canada SOR/98/282 Added section 6.6.1.1 from SOR/98/282 subsection 61.4 |  |
| 8 | 3/13/2024 | Added section 6.6.1.1.5 from SOR/98/282 subsection 61.4 |  |
| 9 | 5/23/2025 | Responsibility Section Clarified: Corrected and clarified the role of MR and/or PRRC to ensure consistent usage throughout the procedure. PMCF Justification Requirement Added: Incorporated language specifying that if Post-Market Clinical Follow-up (PMCF) is not applicable, justification must be documented using Form 9.1-3-3 in accordance with MDR Annex XIV Part B. Class III Clarification: Added a clear statement that Tulip Medical Products currently does not manufacture Class III or implantable devices; removed mandatory submission language for EUDAMED as not applicable. PSUR Submission Note Updated: Clarified that PSURs will be submitted to EUDAMED once the system is fully operational; until then, they will be made available to the Notified Body upon request. Annual Review Requirement Stated: Included explicit instruction in Section 7 (Quality Records) that PMS documentation must be reviewed and updated at least annually or per the PMS Plan. CAPA Traceability Emphasized: Added reference to QSP 8.5-1 to reinforce CAPA linkage and documentation requirements arising from PMS data. Literature Review Scope Refined: Removed IVDR reference from literature evaluation requirements, aligning with Tulip’s CE-marked (non-IVD) portfolio. | Marc Pilkington |
| 10 | 08/21/2025 | 6.6.1.1 added must document the precise start and end dates of the reporting period | Salvador Iniguez |

- PURPOSE -- SCOPE

The purpose of this procedure is to identify and implement a system to proactively collect and review experience gained from devices Tulip Medical Products places on the market, in order to identify any need to apply corrective or preventive actions, and to ensure customer requirements are met.

- Reference DOcuments

US FDA 21 CFR Part 806

US FDA 21 CFR Part 7

Health Canada SOR/98/282

EN ISO 13485

EN ISO 14971

EU Legislation/Guidance

Medical Devices Regulation (EU) 2017/745

CH Legislation/Guidance

(MedDO; SR 812.213)

UK Legislation/Guidance

(UK MDR) 2002

MDR Annex XIV Part B

| QSP 4.2-1 | Document Control |
| --- | --- |
| QSP 5.6-1 | Management Review |
| QSP 5.6-2 | QMS Risk Management Plan |
| QSP 7.3-3 | Risk Management |
| QSP 8.2-2 | Customer Feedback and Complaints |
| QSP 8.5-1 | Corrective and Preventive Action |
| QSP 9.1-0 | MDR Compliance Manual |
| QSP 9.1-2 | Technical Documentation – EU MDR |
| QSP 9.1-4 | Development of Clinical Evaluation Report |
| QSP 9.1-5 | Medical Device Incident Report & Field Safety Corrective Actions - MDR |
| Form 9.1-3-1 | PMS Plan Template |
| Form 9.1-3-2 | PMS Report Template |
| Form 9.1-3-2-CMDR | Health Canada Report Template |
| Form 9.1-3-3 | PMCF Plan Template |
| Form 9.1-3-4 | PMCF Report Template |
| Form 9.1-3-5 | PSUR Report Template |

- DEFINITIONS

Generic device group (MDR Art. 2(7)/IVDR Art. 2(8)): A set of devices having the same or similar intended purposes or a commonality of technology allowing them to be classified in a generic manner not reflecting specific characteristics (also referred to as a “device family”).

Post-market clinical follow-up (PMCF – MDR Annex XIV, Part B, 5): Understood to be a continuous process that updates the clinical evaluation and shall be part of the manufacturer's post-market surveillance plan. When conducting PMCF, the manufacturer shall proactively collect and evaluate clinical data from the use in or on humans of a device which bears the CE marking, and is placed on the market or put into service within its intended purpose as referred to in the relevant conformity assessment procedure, with the aim of confirming the safety and performance throughout the expected lifetime of the device, the continued acceptability of identified risks and to detect emerging risks on the basis of factual evidence.

Post-Market Performance follow-up (PMPF – IVDR Annex XIII, Part B, 4): Understood to be a continuous process that updates the performance evaluation and shall be specifically addressed in the manufacturer's post-market surveillance plan. When conducting PMPF, the manufacturer shall proactively collect and evaluate performance and relevant scientific data from the use of a device which bears the CE marking and is placed on the market or put into service within its intended purpose as referred to in the relevant conformity assessment procedure, with the aim of confirming the safety, performance and scientific validity throughout the expected lifetime of the device, of ensuring the continued acceptability of the benefit-risk ratio and of detecting emerging risks on the basis of factual evidence.

Post-Market Surveillance (PMS – MDR Art. 2 (60)/IVDR Art. 2 (63)): All activities carried out by manufacturers in cooperation with other economic operators to institute and keep up to date a systematic procedure to proactively collect and review experience gained from devices they place on the market, make available on the market or put into service for the purpose of identifying any need to immediately apply any necessary corrective or preventive action.

Post-Market Surveillance (ISO 13485): Systematic process to collect and analyze experience gained from medical devices that have been placed on the market.

- ABBREVIATIONS

AR – Authorized Representative

CA – Competent Authority

CH-REP – Swiss Authorized Representative

EO – Economic Operator

IVDR – In Vitro Diagnostic Devices Regulation (EU)

MDR – Medical Devices Regulation (EU)

MR – Management Representative

NB – Notified Body

PMCF – Post-Market Clinical Follow-up

PMPF – Post-Market Performance Follow-up

PMS – Post Market Surveillance

PMSR – Post Market Summary Report (Canada)

PRRC – Person Responsible for Regulatory Compliance (EU)

PSUR – Periodic Safety Update Report

RMF – Risk Management File

SSCP – Summary of safety and clinical performance

SSP – Summary of safety and performance

UK MDR – Medical Device Regulations 2002

UKRP – UK Responsible Person

- RESPONSIBILITIES

The MR and/or PRRC is responsible for maintaining this procedure in a current state. Use of MR vs. PRRC should be consistent throughout the document, based on role applicability

Marketing or Product Development and the MR and /or PRRC are responsible for creating an annual PMS Plan for Tulip Medical Products’ devices using methodology and analysis of post-market activities.

All employees are responsible for ensuring that any feedback regarding the medical devices provided by the company is passed back to the Regulatory Team in accordance with QSP 8.2-2 Customer Feedback and Complaints.

The Quality Manager or Regulatory Team are responsible for evaluating production and quality data.

The Production Team or Purchasing Department are responsible for evaluating and monitoring Contract Manufacturers’ capability, reliability, quality, etc.

The Regulatory and Sales Teams are responsible for evaluating and monitoring trade shows and articles and publications such as Journals, Periodicals, etc.

The Regulatory Team is responsible for documenting clinical study/performance study, evaluating articles and publications and maintaining the Clinical Evaluation Report.

The Regulatory, Sales, Marketing, and Management teams are responsible for evaluating Customer Surveys or Market Surveys that are performed as part of PMCF MEDDEV 2.12/2 rev.2

- PROCEDURE

  - Overall General Requirements

    - Company must plan, document, implement, and update a PMS system that is proportionate to the risk class and appropriate for the type of device.

    - The PMS system must be an integral part of our QMS.

      - Refer to Flowchart A for PMS Process.

      - Refer to Flowchart B for PMS Interaction with Product Technical Documentation.

    - The PMS system must be planned to actively and systematically gather, record, and analyze pertinent data on the quality, performance and safety of a device through its lifetime to:

      - Obtain the necessary conclusions, and

      - Determine, implement and monitor any CAPA.

  - PMS Sources for Monitoring

    - The processes for collection of PMS data must be documented in the PMS Plan and include, in particular (per MDR/IVDR Annex III, 1.1 (a)):

      - Serious adverse incidents, including information from PSURs and FSN / FSCA,

      - Records referring to non-serious incidents and data on undesirable side-effects,

      - Information from trend reporting,

      - Information from feedback and complaints provided by users, distributors and importers,

      - Relevant market research, literature, articles, journals, databases and/or registers, and

      - Publicly available information about similar devices.

    - Other PMS data may include, as applicable for the device / device family:

      - Information from service and installation activities,

      - Customer surveys or focus groups,

      - Post-Production Risk Reviews,

      - Field evaluations,

      - PMCF MEDDEV 2.12/2 rev.2

      - PMSR SOR-98-282 – subsection 61.4

  - PMS Plan

    - Develop and document a PMS Plan for each device/device family.

    - The PMS Plan shall address the proactive and systemic process to collect the information from the criteria listed in Step 6.2 and per MDR/IVDR Annex III, 1.1 (b), and cover at least the following:

      - Allow for a correct characterization of the performance of the devices, as well as a comparison to be made between the device and similar products available on the market,

      - Determine effective and appropriate methods and processes to assess the collected data,

      - Establish suitable indicators and threshold values used in the continuous reassessment of the benefit-risk analysis and of the risk management for the device,

      - Reference effective and appropriate methods and tools to investigate complaints and analyze market-related experience collected in the field,

      - Specify methods and protocols to manage the events subject to trend reporting, related to any significant increase in frequency or severity of:

- Incidents that are not serious incidents

- Expected undesirable side-effects that could have a significant impact on the benefit-risk analysis, and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits, including the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period;

      - Reference to:

        - Methods and protocols used to communicate effectively with CAs, NBs, EOs and users,

        - Procedures to fulfil the manufacturers obligations documented in the PMS system, PMS Plan and PSUR,

        - Procedures to identify and initiate appropriate measures including corrective actions,

        - Effective tools for tracing and identifying devices for which corrective actions might be necessary.

      - If PMCF is not required, a documented justification using Form 9.1-3-3 must be provided, consistent with MDR Annex XIV, Part B, 5.

    - Use PMS Plan Form to develop the PMS Plan and schedule for each device family. The evaluation of post-market activities is completed at least annually or within the time period established by the PMS Plan.

Note: 	The PMS Plan may vary with each device family depending on the life cycle of the product, e.g., a new device release versus a device that has been on the market for years, sales volume, etc. However, the plan must include the minimum PMS activities as listed in Step 6.2.1. PMS review for an established product may only need to be annual while a newly released product may need to be reviewed more frequently post-launch.  The rationale for frequencies shall be recorded on the PMS Plan.

Note: 	See Table A: PMS & PSUR Reporting

Manufacturers of Class I devices must prepare a PMS report summarizing the results and conclusions of PMS data derived from the PMS Plan together with a rationale and description of any CAPA taken. The report must be updated when necessary and made available to CA upon request.

Manufacturers for Class IIa, IIb and III devices must prepare a PSUR for each device or each category or group of devices summarizing the results and conclusions of PMS data derived from the PMS Plan together with a rationale and description of any CAPA taken. See Table A for reporting timeframes.

  - PMS Data Collection and Review

    - Data is collected in accordance with the PMS Plan and applicable company procedures. The life cycle analysis is defined in the PMS Plan to help determine when trends require action.

Note: 	See Table B for examples of gathering and documenting data.

    - Data is gathered, summarized and used to:

      - Update:

      - benefit-risk determination and risk management files,

      - the design and manufacturing information, instructions for use and labeling,

      - the clinical evaluation report

      - Identify:

      - the need for CAPA or FSN / FSCA,

      - possibilities to improve the device usability, performance and safety,

      - when relevant, how data may contribute to the PMS of other devices, and

      - Detect and report trends per Step 6.3.2.5.

    - Data must be used to update the applicable Technical Documentation.

    - PMS Reports and/or PSURs, including summaries, and trending data and analysis, are reported to the Management Team at the Management Review meeting, reference Management Review Procedure and/or other applicable meetings.

  - Risk Management

    - PMS is a critical component of the overall risk management process and each device’s life cycle in which information is gathered on the identified and unforeseen hazards associated with the finished device.

    - Information gathered as part of this procedure is reviewed continuously against the applicable product / product family Risk Management File (RMF)

    - Address any unidentified or previously unrecognized hazards by updating the RMF.  This may require conducting another risk assessment if a risk evaluation must be completed for the new hazard or risk.

    - New hazards or risks associated with the device may be generated from the following:

      - Complaint where the occurrence has not been observed before and has not been identified in the risk analysis.

      - Feedback on a new indication for use or new feature that may introduce a new unforeseen hazard.

      - A significant increase in the occurrence of a particular complaint or fault condition associated with the medical device.

      - Complaint or feedback where the product has been misused, used abnormally, or not used as intended.

      - Any references, indications of, or new indications of use that may introduce new hazards or risks that are identified in the customer surveys or publications.

      - Any adverse event reports, vigilance reports, product recalls, product corrections, or advisory notices must generate immediate review of the RMF.

    - Enter any new hazards or risks that have been identified on the PMS Report or other applicable documentation.  Ensure the RMF has been reviewed and/or updated as required, including the benefit-risk determination for the device(s).

  - Reporting Results and Initiation of Improvement, Corrective Action or Preventive Action (CAPA)

    - PMS data gathered in accordance with the PMS Plan is documented in the PMS Report, for Class I devices.

      - Canada - Tulip Medical’s regulatory team will prepare in the case of its Class II medical devices, on a biennial basis, a summary report PMSR (use Form 9.1-3-2-CMDR) of the information referred to in subsection 6.6.2.1 that it received or became aware of during the previous 24 months.

              - Form 9.1-3-2-CMDR must document the precise start and end dates of the reporting period.

        - The information to be covered by the summary report is that in respect of

          - adverse effects.

          - reported problems relating to the performance characteristics or safety of the device, including any consumer complaints, received by the manufacturer, importer or distributor after the device was first sold in Canada and all actions taken by the manufacturer, importer or distributor in response to these problems.

          - incidents referred to in SOR-98-282 subsection 59(1); and

          - serious risks of injury to human health that are relevant to the safety of the medical device and are referred to in SOR-98-282 subsection 61.2(2).

        - The summary report shall contain a concise critical analysis of the information referred to in section 6.6.2.1.

        - In preparing the summary report, the holder shall determine, on the basis of the critical analysis, whether what is known about the benefits and risks associated with the medical device has changed in any of the following ways:

          - the potential benefits for patients through the use of the device may be less;

          - in respect of each of the risks,

          - the harm associated with the risk is more likely to occur, or

          - if the harm associated with the risk occurs, the consequences for the health or safety of patients, users or other persons could be more serious; and

          - a new risk has been identified.

        - Include the conclusions reached in Overall Conclusion section of the PMSR.

        - If, in preparing the summary report, the holder concludes that what is known about the benefits and risks associated with the medical device has changed in any of the ways referred to in paragraphs (4)(a) to (c), they shall notify the Minister, in writing, within 72 hours after having reached the conclusion, unless that has already been done

Note:  	For Class IIa, IIb, III device(s) data can be gathered on the PMS Report or via other means for later compilation on PSUR.

      - The results of the PMS data are trended (where applicable), evaluated for product improvements and/or CAPA.

      - Data is compared to quality targets, metrics and/or objectives, as identified in the PMS Plan.

      - When applicable, product improvements and CAPAs are documented and processed in accordance with Document Control Procedure, Design and Development Procedure and/or CAPA Procedure.

    - PMS information must also be incorporated as part of each product’s:

      - RMF, refer to Risk Management Procedure and Step 6.5 for more details,

      - Clinical Evaluation Report, refer to QSP 9.1-4 Development of a CER regarding literature search updates;

      - Technical Documentation and Design History File, when applicable, per QSP 9.1-3 Technical Documentation.

- QUALITY RECORDS

  - The completed PMS Plan, PMS Report and/or PSUR including related supporting data, and applicable records are maintained according to QSP 4.2-1 Control of Records Procedure.

  - Information obtained from marketing surveys, sales contacts, or trade shows shall be maintained with the post-market activities file.

  - All PMS documentation shall be reviewed and updated at least annually, or as specified in the device-specific PMS Plan.

- PROCESS FLOWCHARTS

PMS records retained according to Control of Quality Records procedure

Any issues/ trends?

Initiate as applicable:

- Improvement activities

- Preventive Action

- Corrective Action

Identify PMS sources for monitoring

Develop PMS plan for each device/ device family

Data collected in accordance with PMS plan

PMS data reviewed and trend analysis performed

Results used to update Medical Device File (RM file, CER, etc.) as needed

Yes

No

Continue to monitor PMS plan

Flowchart A: PMS Process

Flowchart B: PMS Interaction with Product Technical Documentation

Technical Documentation

Public Information

Non-vigilance Related Data

Vigilance-related Data

PMCF Studies

PMS Plan

PMCF Evaluation Report

PMS Report / PSUR

Incident Reports, Trend Reports, FSCAs

Periodic Safety Update Report

SSCP

Design History File

Clinical Evaluation Report

Risk Management File

Label - IFU

Table A: PMS & PSUR Reporting

| EUROPE | Regulation References: - MDR; Articles 83, 84, 85, 86, and Annexes II, III |
| --- | --- |
| PMS Reporting – Class I - Prepare a PMS Report summarizing the results and conclusions of PMS data derived from the PMS Plan together with a rationale and description of any CAPA taken. - The PMS Report must be updated when necessary and made available to CA upon request. |  |
| PSUR (Refer to Form 9.1-3-5 PSUR Template) - Class IIa, IIb, III - Prepare a PSUR for each device or each category or group of devices summarizing the results and conclusions of PMS data derived from the PMS Plan. - The PSUR will include a rationale and description of any CAPA taken. - Throughout the life cycle of the device the PSUR must set out the: - Conclusion of the benefit-risk determination, - Main findings of the PMCF, - Volume or sales and an estimate evaluation of the size and other characteristics of the population using the device and. Where practicable, the usage frequency of the device. - Manufacturers of Class IIa devices must update the PSUR when necessary and at least every two (2) years. - The PSUR shall, except for custom-made devices, be part of the Technical Documentation as specified in MDR Annexes II and III. - Manufacturers of Class IIb, III devices must update the PSUR at least annually. - The PSUR shall, except for custom-made devices, be part of the Technical Documentation as specified in Annexes II and III. - For Manufacturers of custom-made devices, the PSUR shall be part of the design, development and manufacturing documentation (refer to MDR Section 2, Annex XIII). |  |

Table B: Examples of Documenting & Gathering PMS Data

Note: Where PMS Report is referenced, revise to include other documentation if compiling for the PSUR and the PMS Report form is not going to be used.

B.1. Customer Feedback - Tulip Medical Products will accept, process, and use feedback from its customers in various ways. Feedback (product problems, issues, etc.) from customers, Field Evaluations, CAs, AEs, etc., are documented and investigated in accordance with the Customer feedback and Complaint Procedure. - Products are labeled with product data (part, serial and/or lot numbers, UDI [as applicable]) and company contact information for traceability purposes, thus facilitating easy communication from the customer to the company on identity of the products. - We sell our products through direct sales, independent sales, or distributors and we maintain a record of all customers who have purchased our products through the sales order process and distribution records for traceability. - Suggestions for improvement and positive feedback are documented in accordance with the Customer feedback and Complaint Procedure. - Customer suggestions should include Marketing, Sales, Product Development and the MR for review. - Document any actions implemented as appropriate with an evaluation regarding customer feedback and complaints. Changes may be made in accordance with Document Control Procedure, Design and Development Prcedure, CAPA Procedure, or other means as applicable. - Alleged product issues or complaints are documented in accordance with the Customer feedback and Complaint Procedure. - Complaints are reviewed by the MR, Regulatory and Management teams to: - Determine if the complaint is valid or requires further investigation. - To determine if CAPA is required to address a significant issue. - Whether a reportable incident has occurred. And if so, report to the appropriate Regulatory Authority in accordance with applicable procedure(s) based on jurisdiction. - The scope of alleged product issues. - The need for market correction or notification for a volatile situation in the market. - Complaints are periodically reviewed and trends are analyzed to determine if additional actions or corrective action is required.

B.2. Country Databases - The MR or designee will perform a search of applicable country-specific databases on at least an annual basis and according to the PMS Plan. - US FDA Database - Determine if any MDRs have been reported for our device or similar devices; www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/search.CFM. Note: The above website may change over time, if the link does not connect go to www.fda.gov, and search for ‘MAUDE’ or ‘Medical Device Reporting.’ - Australian Databases - Determine if any Adverse Event Notifications or Recalls have been reported for our device or similar devices occurred by searching the DAEN or SARA databases http://apps.tga.gov.au/prod/DEVICES/daen-entry.aspx and http://apps.tga.gov.au/PROD/SARA/arn-entry.aspx - EUDEMED once active - Record the search completion on PMS Report and whether or not any action is required to be taken. - Prepare summary of findings (e.g. number per manufacturer, type of report, reason for report) for attachment to the PMS Plan and Summary.

B.3. Production Activities - Our PM or Operations shall monitor production outputs (such as yields, scrap, rework) and report to Operations and Regulatory for review and trending. - Review reasons for scrap and/or rework for trends. Compare data to the process risks to determine if any new risks or increased frequencies are occurring. - Attach trend charts and data, as applicable and record summary on PMS Report. - Review and evaluate negative trends or increased frequencies for CAPA, and if applicable, document and process in accordance Document Control Procedure, Design and Development Prcedure, CAPA Procedure, and/or other means as applicable.

B.4. Field Evaluation or Clinical Studies - Review and evaluate the results of any clinical studies and/or field evaluations, such as market studies, beta-site studies, design validation, etc., for product improvements, product problems and/or CAPAs, and if applicable document and process in accordance with the following procedures: Document Control, Design and Development, Customer feedback and Complaint, and CAPA. - Clinical studies and/or field evaluation reports are generally reviewed during the Design and Development process at applicable design review meetings, reference Design and Development Procedure. - Record a brief summary of any studies in-progress or that have been completed in the PMS-scheduled review. - Reference any protocol, clinical study, and/or report numbers generated and the location of these documents, on PMS Report.

B.5. Articles and Publications - Refer to Development of Clinical Evaluation Report Procedure for evaluation of literature articles related to compliance in the EU for CE-marked products. - Designated personnel, as outlined in PMS Plan, review Journals, Periodicals, Articles, etc., that are relative and/or applicable to Tulip Medical Products’ medical devices, similar devices, or competitors’ products. - Document the Journals, publications, Periodicals, etc., to be reviewed on PMS Plan including the frequency of review, e.g. monthly or quarterly depending on the publication’s issuing cycle. - The MR reviews the complete list of Journals, Periodicals, etc. per the PMS schedule to ascertain if any articles are related to Tulip Medical Products products, our competitors’ products, or products that have a similar intended use. - If the publication contains an article/study, which references a product or product group related to those manufactured by Tulip Medical Products, (whether it references a competitor product, or whether the article is positive or negative towards the product) a copy of the article will be made for inclusion in the PMS file. - As these publications are reviewed, complete PMS Report indicating if the publication has a related article; make a copy of the article for inclusion into the PMS file. Note: Any additional pertinent articles, studies or abstracts reviewed that are not on the schedule will also be documented in applicable area of PMS Report. - Any negative reviews or potential negative reviews are evaluated, and any action taken is documented and processed in accordance with Customer feedback and Complaint Procedure and/or CAPA Procedure.

B.6. Customer and Market Surveys - Customer Surveys are typically handled by the Marketing or Sales group, and are sent to our entire (or a sample of) customer base to gather information on the use of our product. - Customer surveys are generated to ask questions about the product, gather feedback on the use of the product, and help the company understand the acceptance of the product in the customer’s hands. - Review and evaluate the results of any customer surveys for product improvements, product problems and/or CAPAs, and if applicable document and process in accordance with the following procedures: Document Control, Design and Development, Customer feedback and Complaint, and CAPA. - Process any perceived negative feedback, such as low scores on customer surveys or written feedback in accordance with Customer feedback and Complaint Procedure and/or CAPA Procedure. - Document suggestions for improvement on PMS Report and include evaluation or action to be taken. - Periodically review surveys and analyze trends to determine if additional actions or CAPA is required. Document survey results on PMS Report.

B.7. Trade Shows - Individuals that are attending trade shows or conferences where the company’s products are presented should be aware of, and document, any feedback that can be obtained during conversations with customers or potential customers. Note: This can include product improvements, negative comments, complaints, other product problems, customer feedback, additional features, etc. - Document summary of data on PMS Report. - Review results and if applicable process negative / potential negative feedback in accordance with Customer feedback and Complaint Procedure and/or CAPA Procedure.

END OF DOCUMENT

Copyright © 2022 Tulip Medical Products