Tulip Medical

Quality System Procedure

QSP 6.4-1 Cleaning Procedures for Assembly Areas

Approvals

VP Regulatory Affairs	Date

Regulatory Affairs Coordinator	Date

Document Control Administrator	Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 1 | 12/10/2013 | Cleaning procedures updated for all production areas. |
| 2 | 05/07/2014 | Disposable Assembly area updated for new products. |
| 3 | 06/18/2014 | Disposable Assembly Area updated for weekly cleaning of air conditioner filter cover. |
| 4 | 11/13/2014 | Updated Disposable Assembly area cleaning instructions. |
| 5 | 08/30/2018 | Cleaning Procedure updated to match current procedure used. Reformatted. Updates made in response to new requirements outlined in ISO 13485:2016, specifically the inclusion of Section 7.1 Contamination Control. |
| 6 | 5/13/2019 | Rephrase the Purpose, Scope, and Responsibilities to align with the process Changed MR and Operations responsibilities to Production Manager (PM) 4.0 Definition – added Contamination 7.0 Rephrase verbiage 7.3 Removed bullet #8 7.4.1 Revised to read ‘A maximum of four (4) people’ 7.4.2 Revised to read ‘any product or component dropped on the floor should be immediately cleaned until free of particulate contamination. 7.4.9 Revised to read ‘avoid direct contact between shedding cardboard and parts while in the DISPOSABLE ASSEMBLY AREAS’ 7.5.1 Daily Cleaning thru 7.5.6 – reorganized to align with each step 7.5 thru 7.6 – clarify cleaning days 7.6.1 and 7.6.2 – reorganized to align with each step |
| 7 | 8/7/2019 | Updated the following: 1.0 Purpose – added the areas of routine cleaning 7.5.1 Daily Cleaning – added “Clean sticky mat by removing top layer” to 7.5.1.1 7.5.1.3 – corrected the location to “outside the…” 7.5.2 Weekly Cleaning – added (deep cleaning) 7.5.2.1 – added “Clean sticky mat by removing top layer” |
| 8 | 3/31/2020 | 1.0 Purpose and 2.0 Scope – updated 3.0 References – updated 4.0 Definition – included RDC No. 16:2013 5.1.3.4 Contamination Control (Brazil) 5.0 Abbreviation through the entire Process, all sections have been updated to current procedures |
| 9 | 05/10/2023 | - Purpose –Implemented quarterly cleaning and introduced ultrasonic welding in the finished goods area. 3.0 Reference- WI 1 and WI 1.3 have been updated to reflect the current WI name and numbering. 4.0 Definitions- Defined Disposable Assembly Areas and addition of Ultrasonic Welding and finished goods assembly areas. 2.0 Pest Control- Creation of RFD-117 Pest Control Baited Map 4.0 The reference name for the Disposable Production Room has been updated to Disposable Assembly Areas. 5.0 Developed cleaning procedures for all Disposable Assembly Areas, Ultrasonic Welding, and Finished Goods and Receiving Room - Daily, Weekly, and Quarterly. 5.1-3 Added Ultrasonic Welding and Finished Goods Assembly Area Cleaning Procedures for Daily, Weekly, and Quarterly 8.0 Updated cleaning procedures for Reusable- Daily, Weekly, and Quarterly. -Renamed cleaning records for Assembly Areas. 10.Creation of Shipping Area cleaning procedures, Creation of Form 6.4-1-1-5 Sipping Area Cleaning Record |
| 10 | 5/23/2025 | Updated the Scope section to align with Regulation (EU) 2017/745 (MDR) by referencing compliance with General Safety and Performance Requirements (GSPR) related to cleanliness, contamination control, and hygienic manufacturing conditions. Added MDR 2017/745 to the Reference Documents section. |

- PURPOSE

To describe the Daily, Weekly, and Quarterly cleaning of all areas, i.e., Receiving, Shipping, Disposable and Reusable Assembly Areas, Ultra Sonic Welding, and finished goods area, where the product is received and assembled.

- SCOPE

This procedure applies to all personnel involved in the receipt, handling, and assembly of reusable and disposable medical devices. It outlines the cleaning requirements for controlled production areas—including Receiving, Shipping, Disposable and Reusable Assembly Areas, Ultrasonic Welding, and Finished Goods Assembly areas—to prevent contamination that could affect product safety or performance. The procedure ensures that cleaning activities are carried out in a consistent and documented manner, supporting compliance with the General Safety and Performance Requirements (GSPR) related to cleanliness, contamination control, and hygienic manufacturing conditions, as required by Regulation (EU) MDR 2017/745.

- REFERENCE DOCUMENTS

| Regulation (EU) 2017/745 on Medical Devices (MDR) |  |
| --- | --- |
| RDC No. 16:2013 | 5.1.3.4 Contamination Control (Brazil) |
| Form 6.2-1-2 | Training Record |
| Form 6.4-1-1 | Disposable Assembly Area 201 Cleaning Record |
| Form 6.4-1-1-1 | Reusable Assembly Area Cleaning Record |
| Form 6.4-1-1-2 | Receiving Area Cleaning Record |
| Form 6.4-1-1-3 | Disposable Assembly Area 202 Cleaning Record |
| Form 6.4-1-1-4 | Ultra-Sonic Welding and Finished Goods Assembly Area Cleaning Record |
| QSP 4.2-2 | Control of Quality Record |
| QSP 6.2-1 | Training and Qualification |
| RFD-015 | Assembly Areas Gowning Guide |
| RFD-117 | Location of Baited Traps |

- DEFINITIONS

5.1.3.4. Contamination Control. Each manufacturer shall establish and maintain procedures to prevent the contamination of equipment, components, manufacturing materials, intermediates, and finished products by cleaning and disinfecting materials, including hazardous substances or contaminants generated by the production process. A pest control program shall be established, and whenever chemical agents are used, the company shall ensure they do not affect the product quality. (RDC No. 16:2013)

Disposable Assembly Areas – Two assembly areas are utilized to produce non-sterile disposable cannulas and accessories, with all production processes being validated.

The Ultrasonic Welding and Finished Goods Assembly Area is designated for assembling DNFT90G2 and DNTC products and packaging finished goods. It includes pre-assembly tables and a welding machine.

Receiving Area- Incoming product inspection including sampling and receiving in-process and finished sterile/packaged single-use cannula and accessories for Tulip Medical Products.

- ABBREVIATIONS

DAR – Disposable Assembly Room

RR    – Receiving Room

MR   – Management Representative

PM   –  Production Manager

RESPONSIBILITIES

- Document Control Administrator – Maintains and updates all the documents in the Tulip Master File binders,  Departmental binders, and hardcopy folders.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses developing and monitoring the Quality Management System to ensure the accuracy and effectiveness of all Quality System Procedures (QSPs).

The Production Manager The individual assigned to this task is accountable for providing training, supervising, and maintaining records in compliance with this procedure.

Production Personnel - All personnel trained and qualified to perform this procedure are responsible for ensuring compliance. (QSP 6.2-1; Form 6.2-1-2)

PROCEDURE

- Contamination Control

  - Products that are contaminated or potentially contaminated will be immediately removed from production or assembly to prevent contamination of the work environment, personnel, or product.

- Pest Control

  - The area around the building is treated every month with Standard Pest Control.

  - Control of rodents around and within the building are baited traps. No chemicals are used.

  - To find the locations of baited traps, please refer to the RFD-117 Locations baited trap.

- General Information

  - The Production Warehouse floor is cleaned twice weekly with floor cleaning solution and water as part of Tulip Medical's janitorial services. No documentation is required.

  - Health and Clothing Requirements

    - Coughs or sneezes should be directed into your covered arm or shoulder, away from work surfaces, parts, and others. Avoid touching finished devices and packaging.

    - Work clothes worn by production personnel must be clean and free of loose particles such as lint, sweater balls, etc. Heavy sweaters or other “shedding” clothing is not to be worn. Sweatshirts and fleece clothing are allowed, provided it does not have loose particles on them.

.

- General Cleaning Procedures for Disposable Assembly Areas

  - Before performing any cleaning procedures detailed below, employees must adhere to the gowning procedure outlined in the RFD-015 Assembly Areas Gowning Guide.

  - All persons entering the Disposable Assembly Area must comply with the following:

  - No mobile devices, such as cell phones, kindles, etc., are allowed.

- No gum, candy, beverages, or food allowed.

- No make-up, nail polish, or perfumes are allowed.

- No hand jewelry, such as bracelets or rings, to avoid puncturing the gloves.

- Closed-toe and closed-heel shoes MUST be worn.

- Use of tobacco products is not permitted.

- No personnel or visitors with open wounds, cuts, or bandages are allowed unless cleared by the PM.

- All containers, racks, jigs, fixtures, and tools MUST be cleaned to the same level of cleanliness specified for the devices before being brought into the disposable assembly areas.

- Do not store bins on the floor.

  - The finished products assembly area is adjacent to the ultrasonic welding area.

- Disposable Assembly Areas/ Ultrasonic Welding and Finish Goods Area.

  - The following section provides guidelines for cleaning procedures of non-sterile disposable assembly areas. This includes Disposable Assembly Areas 201 and 202 and the Ultrasonic Welding and Finished Goods Assembly Area.

  - Cleaning schedule:

- Daily Cleaning is performed at the beginning of each Assembly Day.

    - Weekly Cleaning is performed every Monday of a Production week.

    - Quarterly Cleaning is performed on the first Production Day of each new Quarter (January, April, July, and October).

Daily Cleaning

    - Use Form 6.4-1-1 and Form 6.4-1-1-3 to initial the daily cleaning procedure for Disposable Assembly areas 201 and 202.

    - Use Isopropyl alcohol wipes to clean the following.

- Glue Station

- Glass Covered and Tables work areas

- Cannulas Curing Trays

- Top of 405 Flood System

- Plastic Mat

- Underside of workstation

    - Daily cleaning of the Ultrasonic Welding and Finished Goods Assembly Area. Use Form 6.1-1-1-4 to Initial the daily cleaning procedure.

    - Use Isopropyl alcohol wipes to clean the following.

- Clean the table that surrounds the Ultrasonic welding machine

- Clean table mats and tables

- Wipe down chair

  - When disinfecting a surface with isopropyl alcohol wipes, use overlapping strokes from top to bottom or left to right.

  - If bins or storage bags are to be used to store in-process products, ensure that these have also been cleaned with a wet nap and disinfected with isopropyl alcohol wipes before being given ample time to dry.

  - Dispose of wipes after use in the trash bins.

Before use, allow these areas to dry. Complete the corresponding room's cleaning record initial task at your station.

- Weekly cleaning

  - Weekly production cleaning occurs every Monday and is supplemented with daily and additional cleaning tasks.

    - For Disposable Assembly areas 201 & 202, Use Form 6.4-1-1 & Form 6.4-1-1-3 to Initial the weekly cleaning procedure.

    - Use Isopropyl alcohol wipes to clean the following.

- Chair seats and backrests (front and back)

- Outsides of cabinets

  - Clean the floor stick mat by removing the top layer

  - Change Trash Can

  - Weekly cleaning of the Ultrasonic Welding and finished goods occurs every Monday and is supplemented with daily and additional cleaning tasks.

    - Use a Swiffer Sweeper Pad to clean the following. Use Form 6.1-1-1-4 to Initial the weekly cleaning procedure.

  - Clean the floors

Before use, allow these areas to dry. Complete the corresponding room's cleaning record initial task at your station.

- Quarterly Cleaning

  - On the first Production Day of each new Quarter (January, April, July, and October), perform the daily cleaning procedures and additional cleaning duties for Assembly Areas 201 and 202 and the Ultrasonic Welding and Finished Goods Assembly Area.

  - Clean the air conditioner vent with a Swiffer Sweeper pad. Visually inspect the vent cover for dust/dirt.

  - To clean the floor first, sweep it with a broom. Then, use a Swiffer pad to clean the floor thoroughly.

  - Wipe walls with an Isopropyl alcohol wipe.

Before use, allow these areas to dry. Complete the corresponding room's cleaning record initial task at your station.

- Reusable Assembly Area Cleaning Procedure
The daily, weekly, and quarterly cleaning procedures are outlined below, with additional instructions in sections 8.1.2 to 8.3.

  - Cleaning Schedule

- Daily Cleaning is performed at the beginning of each Assembly Day.

    - Weekly Cleaning is performed every Monday of a Production week.

    - Quarterly Cleaning is performed on the first Production Day of each new Quarter (January, April, July, and October).

    - To sign off on the cleaning procedure for the Reusable Assembly Area, please use Form 6.4-1-1-1.

    - Daily Cleaning

    - Use Isopropyl alcohol wipes to clean the following.

- Glass-covered work areas

- Cannula curing trays

- Cabinet handles and Inside and Outside door handles

  - Weekly cleaning is performed every Monday during production weeks, along with daily procedures and additional cleaning tasks.

    - Use Isopropyl alcohol wipes to clean the following.

- Chair seats and backrests

- The top and sides of the refrigerator used to store the glue

- Tops of cabinets and storage bins

    - Sweep the tops of workbenches inside the assembly area to remove dust and debris, including where stems are cut and polished.

    - Use the Swiffer Sweeper to clean the floors of the Assembly Room.

    - Allow these areas to dry before use.

  - Quarterly Cleaning is performed on the first Production Day of each new Quarter (January, April, July, and October). Perform the daily cleaning procedures as per routine and additional cleaning tasks listed below.

    - Clean the top of the Curing Oven.

    - Sweep the tops of worktables outside the Production Room to remove dust and debris, including where stems are cut and polished.

    - Use Isopropyl alcohol wipes to wipe down the walls of the Production Area.

    - Use the Swiffer Sweeper to clean the floors of the Production Area.

    - Allow all areas to dry before use.

    - Empty all trash bins.

- Receiving Area Cleaning

  - The following procedures are done when the product is delivered and ready for inspection. (To sign off on the cleaning procedure for the Receiving area, please use the Form 6.4-1-1-2 )

  - Prepare the Receiving Area for the inspection of products

    - Daily Cleaning

    - Use Isopropyl alcohol wipes to clean the following.

  - Clean bins that will be used to store inspected product

  - Clean counters

    - Weekly cleaning is performed every Monday during production weeks, along with daily procedures and additional cleaning tasks.

- Wipe down the scale.

- Underside of work area and seats

- Empty all trash bins

    - Quarterly Cleaning is performed on the first Production Day of each new Quarter (January, April, July, and October). Perform the daily cleaning procedures as per routine and additional cleaning tasks listed below.

    - Use a Swiffer Sweeper pad to wipe down the walls and ceiling of the Receiving Area.

  - Use the Swiffer Sweeper to clean the floors of the Receiving Area.

- Shipping Department

  - Before starting the shipping process, the following cleaning procedures must be followed.

To sign off on the cleaning procedure for the shipping area, please use the Form 6.4-1-1-5 Shipping Area Cleaning Record.

  - Daily Cleaning

    - Use Isopropyl alcohol wipes to clean the following.

- Wipe down the Tables

- Wipe down the weight scale

  - Weekly Cleaning

    - Weekly cleaning is performed every Monday during production weeks, along with daily procedures and additional cleaning tasks listed below.

  - Clean Transport Carts

  - Quarterly Cleaning is performed on the first Production Day of each new Quarter (January, April, July, and October). Perform the daily cleaning procedures as per routine and additional cleaning tasks listed below.

  - Sweep under box racks and surroundings.

  - Mop using a Swiffer Sweeper Pad

  - Clean Air Vent