Tulip Medical

Quality System Procedure

QSP 7.4-5:  Process Validation

Approvals

VP Regulatory Affairs	Date

Regulatory Affairs Coordinator	Date

Document Control Administrator	Date

Revision History

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 |  | Initial release. |
| 1 |  | Added process validation requirements for state-of-the art standards; changed font to Calibri and size to 10 |
| 2 | 6/24/2020 | 6.0 Procedure, added 6.2.7.6 Software Validation |
| 3 | 5/23/2025 | Aligned QSP 7.4-5 with EU MDR 2017/745 requirements. Added reference to Article 10(9) and Annex IX, Chapter I, Section 3.3 to the Purpose section, establishing regulatory basis for process validation where outputs cannot be verified post-production. In the Process Validation Master Plan section, clarified linkage to GSPR under Annex I for biocompatibility and sterility standards. Added reference in the Final Report section to MDR Annex IX, Section 3.3(c) requiring documentation of validation status and corrective actions. In Maintaining a State of Validation, added citation of Article 10(9)(e) emphasizing ongoing process monitoring and validation controls. Updated References section to include applicable MDR clauses (Article 10(9), Annex I, and Annex IX). |

# Purpose

This procedure defines process validation methods and the processes involved in developing a Process Validation Master Plan.  This procedure is established to support compliance with the requirements of Article 10(9) and Annex IX, Chapter I, Section 3.3 of Regulation (EU) 2017/745 (MDR), which mandate validation of manufacturing processes that cannot be fully verified by subsequent monitoring or measurement.

# Scope

This procedure applies to all equipment and processes used for product development at Tulip Medical Products.

# References

QSP 7.3-1 Design and Development

FRME-0711 Equipment Installation Qualification Checklist

ISO 14971 Application of risk management to medical devices

ISO 13485 Quality management systems: Requirements for regulatory purposes

ISO 10993 Biological evaluation of medical devices QSP 8.5-1 Corrective and Preventive Action

QSP 4.2-1 Document Control

QSP 4.2-2 Control of Quality Records

Regulation (EU) 2017/745 on Medical Devices (MDR), Article 10(9)

Regulation (EU) 2017/745, Annex I – General Safety and Performance Requirements (GSPR)

Regulation (EU) 2017/745, Annex IX, Chapter I – Quality Management System Requirements

# Definitions

Installation Qualification (IQ) – documentation that equipment has been appropriately installed and establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification and that the recommendations of the supplier of the equipment are suitably considered.  Equipment will be entered into the calibration and preventative maintenance program.

Operation Qualification (OQ) – establishing by objective evidence the nominal operating ranges for each product processed by the equipment.

Performance Qualification (PQ) – establishing by objective evidence that the equipment, under anticipated conditions, consistently produces a product which meets all predetermined requirements.

Quality Check (QC) – a check by a trained worker to ensure that the product meets all predetermined requirements.  QCs will be recorded on production forms for receipt inspection.

Process Validation (PV) – establishing by objective evidence that the process consistently produces a result or product meeting its predetermined requirements.

Process Validation Protocol - a document stating how validation will be conducted, including test parameters, product characteristics, manufacturing equipment, and decision points on what constitutes acceptable test results.

Validate – to establish documented evidence that provides a high degree of assurance that a manufacturing process or test method will consistently produce a product or test result meeting its pre-determined specifications and quality attributes throughout a specified operating range.

Process Validation Master Plan (PVMP) – a living document which identifies validation rationales, activities, deliverables and responsibilities.

State of the Art – developed stage of technical capability at a given time as regards products, processes, and services, based on the relevant consolidated findings of science, technology and experience.

Verification – confirmation by examination and provision of objective evidence that the specified requirements have been fulfilled.

# Responsibilities

Production Manager is responsible for carrying out Installation Qualifications, Operational Qualifications, Performance Qualifications and Process Validations and to document the qualifications and validation in a written protocol and report.

It is the responsibility of the Management Representative to review and approve the Process Validation Master Plan and all qualification/validation protocol and test reports.

# Procedure

This process validation procedure is developed in accordance with MDR 2017/745 Annex IX and Article 10(9), which require that for processes where the output cannot be verified by subsequent monitoring or measurement, process validation must be performed.

## Verify or Validate?

Determine if the process should be verified or validated by using the process validation decision tree (Figure 1) below:

START: Can the process output be verified by subsequent monitoring or measurement?

Is verification alone sufficient to eliminate unacceptable risk & cost effective?

Verify and control the process.

Validate

Yes

No

Yes

No

Redesign Product and/or Process

Choose

Figure 1: Process Validation Decision Tree

## Process Validation Activities

### Process Validation (PV) Breakdown

Process validation encompasses those activities carried out for Installation Qualification (IQ), Operational Qualification (OQ) and Performance Qualification (PQ) within a process.

Once the above qualifications are complete, the process validation for an entire process can commence using a statistically valid sampling plan.  Refer to Figure 2for an example of process validation activities.  Process validation is accomplished when 100% of the specified number of samples/lots meets all required specifications.

Process Validation

Assembly Step 4

Send to Eagle

Package

(IQ, OQ, PQ)

Sterilize

(PV)

Product Distribution

QC

QC

QC

Raw Material

Assembly Step 1

(IQ, OQ, PQ)

QC

Assembly Step 2

Assembly Step 3

(IQ, OQ, PQ)

QC

Figure 2: Process Validation: Example of a process flowchart with several levels of qualification

### Process Validation Master Plan

A Process Validation Master Plan (PVMP) is a document created during the early stages of product development.  The Process Validation Master Plan is updated when design changes are incorporated into the process and/or product.

The plan shall involve those functions involved in Product Development and may include representatives from personnel with expertise in Quality Assurance, R&D, Manufacturing and Regulatory Affairs.

The Process Validation Master Plan identifies those processes to be validated.  For multiple tests a Quality Plan will be used for scheduling and tracking completion.

The Process Validation Master Plan additionally provides the following information:

Rationales for the validation approach

Activities involved in the process validation (IQ, OQ, PQ, etc.)

Deliverables: defining the boundaries of product compliance

State of the art standards to be used for biocompatibility/sterility testing (eg. ISO 13485, 14971, 10993, etc)

- Note: These standards are applied in accordance with MDR 2017/745, Annex I – General Safety and Performance Requirements (GSPR)

Responsibilities

The Process Validation Master Plan shall be documented and approved according to the QSP 4.2-1 Document Control procedure before any elements of a validation is conducted.   Form 5.4-1-1 (Quality Plan) may be used for simple processes/products.

The Process Validation Master Plan is reviewed periodically and updated as appropriate.

### Process Validation Protocol

Separate protocols may be developed for each phase of process validation with the exception for Installation Qualification (IQ).  Installation qualification does not require a protocol and report except for the Installation Qualification form (Form 7.4-5-1).  IQs will be used on all new equipment installation.

- Process validation protocols will be documented on the Test Protocol Templates and Test Report Templates in use.

### Installation Qualification (IQ)

- Installation qualifications are carried out for all equipment to ensure equipment is installed and is working correctly.  Form 7.4-5-1 Installation Qualification will be used in place of a written protocol and report.

### Operational Qualification (OQ)

Once product requirements have been established and equipment installation qualification activities are complete, operational qualification activities can be conducted.  Each product processed with the equipment requires an OQ as the nominal operating range of the equipment settings may vary between products.

Qualification activities at the equipment vendor’s site prior to equipment shipment shall not replace equipment qualification.  However, copies of the vendor’s qualification studies may be used as a guide to supplement equipment qualification.

Prepare a Test Protocol Template.

In the protocol define process limits and specification limits.

Define the operating control limits (time, temperature, pressure, etc.)  Process parameters should be challenged to assure that they will result in a product or process that meets all defined requirements and does not fail under anticipated operating conditions.

Establish operating procedures and/or work instructions, as appropriate.

Perform testing as outlined in the protocol ensuring:

personnel are properly trained prior to testing being carried out

samples are representative in a truly manufacturing environment (there is no set number of samples that need to be determine during OQ)

Generate a report once testing is complete and revise procedures/work instructions as required.

### Performance Qualification (PQ)

The objective of performance qualification is to demonstrate the process will consistently produce acceptable product under anticipated conditions.

Prepare a Test Protocol Template.

Performance qualification considerations include:

Actual product and process parameters and procedures established in OQ

Acceptability of the product

Assurance of process capability as established in OQ

Process repeatability, long term process stability

Establish QC as required.

Challenges to the process should simulate conditions that will be encountered during actual manufacturing.  Challenges should include the range of conditions as defined by the various action levels as established in the OQ phase.  The challenges should be repeated enough times to assure that the results are meaningful and consistent.

Execute and monitor the Process Qualification run per the PQ protocol ensuring:

personnel are properly trained prior to testing being carried out

samples are representative in a truly manufacturing environment (there is no set number of samples that need to be determined during PQ).

No changes are allowed in process parameters or PQ inspection methods during the PQ runs.  If changes are necessary, amend the protocol with statement for reason for change and justification for change.  Repeat the testing as needed.

Create a Test Report once testing is complete.

### Process Validation (PV)

Process validation testing for an entire process will commence once all IQ, OQ and PQ’s have been completed.

Determine a set number of units for process validation using statistically valid techniques.  There should be a scientific rational for the number of units in the protocol.  Financial constraints will also be factored in.

Perform testing at nominal operating specifications per given state of the art standard(s).

Verify 100% of finished product for validation.

Utilize statistical methods to show process capability

Process validation is achieved when you can prove you can consistently meet product specs with successful quality control (no lot failures).

Software Process Validation – only applies to software that affects product and the quality system.

        - Software process shall be validated prior to use.

        - Software shall be validated to the intended use of the process.

        - All validations and protocols shall be documented.

### Final Report

At the conclusion of validation activities, a final report shall be prepared. This report should summarize and reference all protocols and results.  It should derive conclusions regarding the validation status of the process.  The final report shall also document conformity with MDR 2017/745 Annex IX, Chapter I, Section 3.3(c), ensuring that results of validation meet pre-established acceptance criteria and that any deviations are addressed through appropriate corrective actions. This report will be filed in the Design History Folder.

### Maintaining a State of Validation

Trends in the process should be monitored to ensure the process remains within the established parameters.

When monitored data on quality characteristics demonstrates a negative trend, the cause should be investigated.  Corrective actions should be taken according to QSP 8.5-1 Corrective and Preventive Action and revalidation may be considered. Ongoing validation monitoring shall be performed in line with MDR Article 10(9)(e), requiring manufacturers to apply quality assurance measures to ensure continued conformity of the device.

In maintaining a state of validation, the process validation protocol will be reviewed during any changes to the process.  QSP/WIs will be changed as necessary to support.  The Process Validation Final report will be amended as necessary based on these changes with additional reports as needed.   The results of the review shall be documented as an annex to the report.  The review should include a statement that as a result of the review the process is still validated and is in a state of control.

### Revalidation

When major changes or process deviations occur, the process shall be reviewed and evaluated.  Perform revalidation where appropriate.

Examples of reasons for revalidation:

Change(s) in process that may affect quality (i.e. procedures, equipment, etc.)

Investigation of negative trend indicators

Change in product design effecting process

Process transfer from one room or facility to another

Change in scope of process application

Revalidation may not need to be as extensive as the initial validation if the situation does not require that all aspects of the original validation be repeated.  For example, if new equipment is purchased for a validated process, the IQ portion will need to be repeated.  However, some of the OQ aspects have already been established.  Some elements of PQ may need to be repeated, depending on the impact of the new equipment.

# Records

The following records are official outputs of this Operating Procedure:

Form # 7.4-5-1 Installation Qualificaiton,

Test Protocols,

Test Reports,

Process Validation Master Plan

- These records will be controlled/retained per QSP 4.2-2 Control of Quality Records.