Tulip Medical

Quality System Procedure

QSP 9.1-6-1

Unique Device Identification System

Approvals

V.P Regulatory Affairs 						             Date

Regulatory Affairs Coordinator					             Date

Document Control Administrator					              Date

Revision History

| Rev | CR# | Effective / Approval Date | Description of Change | Originator |
| --- | --- | --- | --- | --- |
| 0 |  | 5/23/2025 | Release documentation - - Implementation of document for EU Medical Device Regulations (MDR) 2017/745 Release of form: Form 9.1-6-1 | Marc Pilkington |

- PURPOSE - SCOPE

The purpose of this procedure is to establish the process for implementing a system for the Unique Device Identifier (UDI) to identify Tulip Medical Product’s devices through distribution and use.  Medical devices, other than custom-made, investigational devices or devices for performance study, must be properly identified and include a label showing the plain-text version and automatic identification and data capture (bar code) of the assigned UDI number.

The procedure defines the requirements for submitting the product information to be entered in the European Union’s (EU) database on medical devices (“Eudamed”) and assure this information is properly maintained.  Any new products offered by the company must be considered for the scope of this procedure and assignment of a UDI number.

Note: This document, at time of release, contains the current requirements for the EU MDR; updates may be required for any new generated guidance, Delegating and Implementing Acts, Common Specifications, Harmonized standards, etc.

Note: MDR Art. 27 discusses UDI System and Art. 28 discusses UDI Database, which has yet to be established at time of release of this procedure.

Note: MDR Art. 123 (3):

(f) for implantable devices and for class III devices Article 27(4) shall apply from 26 May 2021. For class IIa and class IIb devices Article 27(4) shall apply from 26 May 2023. For class I devices Article 27(4) shall apply from 26 May 2025;

(g) for reusable devices that shall bear the UDI carrier on the device itself, Article 27(4) shall apply from two years after the date referred to in point (f) of this paragraph for the respective class of devices in that point.

- Reference DOcuments

  - Medical Device Regulation (MDR) 2017/745

  - Eudamed, Directive 95/46/EC of the European Parliament and of the Council

  - IMDRF/UDI WG/N7FINAL: 2013 UDI Guidance: Unique Device Identification (UDI) of Medical Devices 18 Dec. 2013 www.imdrf.org/docs/imdrf/final/technical/imdrf-tech-131209-udi-guidance-140901.pdf

  - 2 IMDRF/UDI WG/N48 FINAL: 2019 Unique Device Identification system (UDI system) Application Guide  www.imdrf.org/docs/imdrf/final/procedural/imdrf-proc-190321-udide-n53.pdf

  - Unique Device Identification (UDI) System under the EU Medical Device Regulations 2017/745 and 2017/746 https://ec.europa.eu/docsroom/documents/36664

  - MDCG Documents

| Reference | Title | Publication |
| --- | --- | --- |
| MDCG 2018-1 v3 | Guidance on basic UDI-DI and changes to UDI-DI | March 2020 |
| MDCG 2019-1 | MDCG guiding principles for issuing entities rules on basic UDI-DI | January 2019 |
| MDCG 2019-2 | Guidance on application of UDI rules to device-part of products referred to in article 1(8), 1(9) and 1(10) of Regulation 745/2017 | February 2019 |
| MDCG 2018-2 | Future EU medical device nomenclature - Description of requirements | March 2018 |
| MDCG 2018-3 | Guidance on UDI for systems and procedure packs | October 2018 |
| MDCG 2018-4 | Definitions/descriptions and formats of the UDI core elements for systems or procedure packs | October 2018 |
| MDCG 2018-5 | UDI assignment to medical device software | October 2018 |
| MDCG 2018-6 | Clarifications of UDI related responsibilities in relation to article 16 | October 2018 |
| MDCG 2018-7 | Provisional considerations regarding language issues associated with the UDI database | October 2018 |
| Reference | Title | Publication |
| MDCG 2018-1 v3 | Guidance on basic UDI-DI and changes to UDI-DI | March 2020 |

  - QSP 4.2-1 Document Control

  - QSP 9.1-1-1 MDR Compliance Manual

  - QSP 9.1-3 Technical Documentation – MDR

  - Form 9.1-6-1 EU UDI Number Listing Template

- DEFINITIONS

  - Automatic Identification and Data Capture (AIDC) (Annex VI Ch C): a technology used to automatically capture data. ADIC technologies include bar codes, smart cards, biometrics and RFID.

  - Basic UDI-DI (Annex VI Ch C): the primary identifier of a device model. It is the DI assigned at the level of the device unit of use. It is the main key for records in the UDI database and is referenced in relevant certificates and EU declarations of conformity.

  - Configurable Device (Annex VI Ch C): a device that consists of several components which can be assembled by the manufacturer in multiple configurations. Those individual components may be devices in themselves.

  - Configuration (Annex VI Ch C): a combination of items of equipment, as specified by the manufacturer, that operate together as a device to achieve an intended purpose. The combination of items may be modified, adjusted or customized to meet specific needs.

  - Device Package: a package that contains a fixed quantity of a particular version of a model of a device.

  - Eudamed: the European database on medical devices (Eudamed) that integrates different electronic systems to collate and process information regarding devices on the market and the relevant economic operators (EO), certain aspects of conformity assessment, notified bodies, certificates, clinical investigations, vigilance and market surveillance. {Currently delayed until May 2022}

  - Human Readable Interpretation (HRI) (Annex VI Ch C): a legible interpretation of the data characters encoded in the UDI carrier.

  - Issuing Entities: an organization designated by the Commission to operate a system for the issuance of UDIs.

  - Label (MDR Art. 2 (13)): the written, printed or graphic information appearing on either the device itself, or on the packaging of each unit or on the packaging of multiple devices.

  - Lot number: a distinctive combination of numbers, from which the history of the manufacturing, packaging, labeling and distribution of a lot or batch of finished devices can be determined.

  - Packaging Levels: the various levels of device packaging that contain a defined quantity of devices, such as a carton or case.

  - Radio Frequency Identification (RFID) (Annex VI Ch C): a technology that uses communication through the use of radio waves to exchange data between a reader and an electronic tag attached to an object, for the purpose of identification.

  - Serial Number: a distinctive combination of letters/numbers, from which the history of the manufacturing, packaging, labeling and distribution of a specific device / unit can be determined and traced.

  - Shipping Containers (Annex VI Ch C):  a container in relation to which traceability is controlled by a process specific to logistics systems.

  - Single Registration Number (SRN): obtained by the competent authority from the electronic system and issued to the manufacturer, the authorized representative (AR), the importer or the System and Procedure Pack Producer (SPPP).

  - System and Procedure Packs (SPP):

    - Procedure Pack (MDR Art. 2 (10)): a combination of products packaged together and placed on the market with the purpose of being used for a specific medical purpose;

    - System (MDR Art. 2 (11)): a combination of products, either packaged together or not, which are intended to be interconnected or combined to achieve a specific medical purpose.

  - Technical Documentation (TD): the name given to the EU regulatory documentation / file for all medical devices.

  - Unique Device Identifier (UDI) (MDR Art. 2(15)): a series of numeric or alphanumeric characters that is created through a globally accepted device identification and coding standard and that allows unambiguous identification of specific devices on the market. The UDI is comprised of the UDI-DI and the UDI-PI.

    - UDI-DI (Annex VI Ch C): a unique numeric or alphanumeric code specific to a model of device and that is also used as the ‘access key’ to information stored in a UDI database.

    - UDI-PI (Annex VI Ch C): a numeric or alphanumeric code that identifies the unit of device production. The different types of UDI-PIs include serial number, lot number, software identification and manufacturing or expiry date or both types of date.

    - UDI Carrier (Annex VI Ch C): the means of conveying the UDI by using AIDC and, if applicable, its HRI. UDI carriers include, among others, ID/linear bar code, 2D/Matrix bar code, RFID.

    - Unit of Use DI (Annex VI Ch C): serves to associate the use of a device with a patient in instances in which a UDI is not labelled on the individual device at the label of its unit of use; for example in the event of several units of the same device being packaged together.

  - Version or Model: all devices that have specifications, performance, size, and composition, within limits set by the manufacturer.

- ABBREVIATIONS

  - AIDC – Automatic Identification and Data Capture

  - AR - Authorized Representative

  - CR – Change Request

  - DEHP - Di(2-ethylhexyl)phthalate

  - DHR – Device History Record

  - EO - Economic Operators

  - EU – European Union

  - FSCA - Field Safety Corrective Action

  - HRI – Human Readable Interpretation

  - PRRC – Person Responsible for Regulatory Compliance (EU)

  - RFID - Radio Frequency Identification

  - SPP - System and Procedure Packs

  - SPPP - System and Procedure Pack Producer

  - SRN – Single Registration Number

  - TD- Technical Documentation

  - UDI – Unique Device Identifier

  - UDI-DI – Unique Device Identifier + Device Identifier

  - UDI-PI - Unique Device Identifier + Product Identifier

- RESPONSIBILITIES

  - The PRRC is responsible for maintaining this procedure in a current state as well as monitoring the activities associated with UDI assignment, application, and removal.

  - PRRC or designee is responsible for obtaining a Basic UDI-DI from an issuing entity and assigning to the device(s), including a system or procedure pack and providing it to the EU UDI database.

    - Before placing a device on the market, other than custom-made, investigational devices, or devices for a performance study, the PRRC or designee is responsible to enter or if, already provided, verify in GS1 and Eudamed the information is correct and shall thereafter keep the information updated.

  - PRRC and Product Development are responsible for:

    - Ensuring that a human readable form and AIDC form of the UDI is located on the finished device and / or finished product labeling, as applicable;

    - The correct labeling information and any updates to existing version or models of devices; and

    - Considering any new versions or models for the UDI number.

  - Sales, Marketing and Customer Service & Production are responsible for ensuring they are aware of the product offerings and UDI information for any customer calls that involve dissatisfaction, complaints, or potential adverse events.

- PROCEDURE

  - General Requirements

    - Tulip Medical Products shall ensure that all device product labeling shall contains the UDI. The UDI does not replace any other marking or labelling requirements, as identified in Annex 1 of the MDR.

    - The manufacturer shall assign and maintain unique UDIs for its devices using GS1.

    - Only the manufacturer may place the UDI on the device or its packaging.

    - Tulip Medical Products shall ensure only coding standards provided by issuing entities designated by the Commission are used.

    - Tulip Medical Products will ensure the UDI is used for reporting serious incidents and FSCAs.

    - The Basic UDI-DI of the device(s) will appear in the TD and on the EU Declaration of Conformity.

    - Tulip Medical Products shall ensure a current list of all UDIs, that it has assigned, is included in all TD.

    - Tulip Medical Products will update the UDI database record within 30 days of a change being made to an element, which does not require a new UDI-DI. Refer to Step 6.7 for Changes to Devices.

    - Relevant obligations arising from MDR Art. 16(1);

      - Any distributor, importer or other natural or legal person that assumes the obligations mandatory on manufacturers in accordance with Art. 16(1), assumes all the relevant responsibilities related to UDI, including UDI labelling.

      - This means that those EO(s) must also apply for registration as Manufacturers, receive a SRN, apply for the appropriate conformity assessment procedure and provide UDI-product registration.

      - However, in accordance with the provision of Art. 16(1)(a), when a distributor or importer enters into an agreement with a manufacturer whereby the manufacturer is identified as such on the label, the manufacturer is responsible for meeting the requirements placed on manufacturers in this Regulation, including the relevant UDI obligations.

  - UDI Requirements

    - A UDI shall be assigned to the device itself or its packaging. Higher levels of packaging shall have their own UDI.

    - Shipping containers are exempted. For example, a UDI shall not be required on a logistics unit; where a healthcare provider orders multiple devices using the UDI or model number of individual devices and the manufacturer places those devices in a container for shipping or to protect the individually packaged devices, the container (logistics unit) shall not be subject to UDI requirements.

    - The UDI shall contain two parts, a UDI-DI and a UDI-PI.

    - The UDI-DI shall be unique at each level of device packaging.

    - If a lot number, serial number, software identification or expiry date appears on the label, it shall be part of the UDI-PI. If there is also a manufacturing date on the label, it does not need to be included in the UDI-PI. If there is only a manufacturing date on the label, this shall be used as the UDI-PI.

    - Each component that is considered to be a device and is commercially available on its own shall be assigned a separate UDI unless the components are part of a configurable device that is marked with its own UDI.

    - Systems and procedure packs (per MDR Art. 22) shall be assigned and bear their own UDI.

    - The manufacturer shall assign the UDI to a device following the relevant coding standard from the issuing entity.

    - A new UDI-DI shall be required whenever there is a change that could lead to misidentification of the device and/or ambiguity in its traceability. Refer to Step 6.9 Changes to Devices.

    - If the [COMPANY / MANUFACTURER] repackages and/or relabels devices, with their own label they shall retain a record of the original device manufacturer's UDI.

    - The UDI number is used to identify only one version, model, or part number of a device.  Even though there may be multiple devices in a product family, only one unique UDI number is assigned per model or part number of a device.

    - When future devices are added they shall be evaluated for UDI requirements as part of the design and development process.

  - UDI Carrier

    - Tulip Medical Products shall ensure the UDI carrier (AIDC and HRI presentation of the UDI) is placed on the label of the device and all higher levels of packaging. Note that higher levels do not include shipping containers.

      - The UDI carrier may be placed on the next higher packaging level if there are significant space constraints.

    - For single-use devices of classes I and IIa packaged and labelled individually, the UDI carrier shall not be required to appear on the packaging, but it shall appear on a higher level of packaging, e.g. a carton containing several individually packaged devices.

      - However, when the healthcare provider is not expected to have access, in cases such as in home healthcare settings, to the higher level of device packaging, the UDI shall be placed on the packaging of the individual device.

    - For devices exclusively intended for retail point of sale the UDI-PIs in AIDC shall not be required to appear on the point of sale packaging.

    - When AIDC carriers other than the UDI carrier are part of the product labelling, the UDI carrier shall be readily identifiable.

    - If linear bar codes are used, the UDI-DI and UDI-PI may be concatenated or non-concatenated in two or more bar codes. All parts and elements of the linear bar code shall be distinguishable and identifiable.

    - If there are significant constraints limiting the use of both AIDC and HRI on the label, only the AIDC format shall be required to appear on the label.

      - For devices intended to be used outside healthcare facilities, such as devices for home care, the HRI shall however appear on the label even if this results in there being no space for the AIDC.

    - The HRI format shall follow the rules of the UDI code-issuing entity.

    - When using RFID technology, a linear or 2D bar code in line with the standard provided by the issuing entities shall also be provided on the label.

    - Devices that are reusable shall bear a UDI carrier on the device itself.

      - The UDI carrier for reusable devices that require cleaning, disinfection, sterilization or refurbishing between patient uses shall be permanent and readable after each process performed to make the device ready for the subsequent use throughout the intended lifetime of the device.

      - The above requirement shall not apply to devices in the following circumstances:

        - Any type of direct marking would interfere with the safety or performance of the device;

        - The device cannot be directly marked because it is not technologically feasible.

    - The UDI carrier shall be readable during normal use and throughout the intended lifetime of the device.

    - If the UDI carrier is readily readable or, in the case of AIDC, scannable, through the device's packaging, the placing of the UDI carrier on the packaging shall not be required.

    - In the case of single finished devices made up of multiple parts that must be assembled before their first use, it shall be sufficient to place the UDI carrier on only one part of each device.

    - The UDI carrier shall be placed in a manner such that the AIDC can be accessed during normal operation or storage.

    - Bar code carriers that include both a UDI-DI and a UDI-PI may also include essential data for the device to operate or other data.

  - UDI Database

    - Tulip Medical Products is responsible for the initial UDI submission in the UDI database, as well as any updates to the data or device elements.

      - The data for new UDI-DIs shall be available at the time the device, other than a custom-made device, investigational device, or device for performance study, is placed on the market.

      - Internationally-accepted standards for data submission and updates shall, wherever possible, be used by the UDI database.

Note: The presence of the device UDI-DI in the UDI database shall not be assumed to mean that the device is in conformity.

    - Obtain Form 9.1-6-1 (EU UDI Number Listing Template) and record the issuing entity and the data in Step 6.4.3., prior to entering into the UDI Database.

      - When UDI-DI obtained record on Form 9.1-6-1.

Note: All of the company’s products that are being assigned a UDI number need to be entered on the EU UDI Number Listing, Form 9.1-6-1.

Note: The EU UDI Number Listing can be maintained in electronic format or as an electronic database; assure that the information listed has a revision date.

    - Tulip Medical Products shall provide to the UDI database the following device data elements in Part B of Annex VI:

- quantity per package configuration;

- the Basic UDI-DI and any additional UDI-DI’s;

- device production controls; e.g. expiry date or manufacturing date, lot number, serial number;

- unit of use UDI-DI, if applicable (where a UDI is not labelled on the device at the level of its unit of use, a ‘unit of use’ DI shall be assigned so as to associate the use of a device with a patient);

- name and address of the manufacturer, including SPPP, as indicated on the label;

- the SRN issued;

- the name and address of the AR as indicated on the label, if applicable;

- the medical device nomenclature code, per MDR Art.26;

- risk class of the device;

- name or trade name, if applicable;

- device model, reference or catalog number, if applicable;

- clinical size (including volume, length, gauge, diameter), if applicable;

- additional product description (optional)

- storage and/or handling conditions, if applicable (as indicated on the label or in the instructions for use);

- additional trade names of the device, if applicable;

- labelled as a single-use device (y/n);

- the maximum number of reuses, if applicable;

- device labelled sterile (y/n);

- need for sterilization before use (y/n);

- containing latex (y/n);

- information labelled in accordance with Annex I, Section 10.4.5, where applicable (substances which are carcinogenic, mutagenic or toxic to reproduction; substances having endocrine-disrupting properties);

- URL for additional information, such as electronic instructions for use (optional);

- critical warnings or contraindications, if applicable;

- status of the device (on the market, no longer placed on the market, recalled, FSCA initiated).

    - In addition to above the following, if applicable, from Annex VI, Part A:

      - reprocessed single-use device (y/n);

      - in case of devices listed in Annex XVI, specification as to whether the intended purpose of the device is other than a medical device;

      - in the case of devices designed and manufactured by another legal or natural person as referred to in MDR Art. 10(15), the name, address and contact details of that legal or natural person.

    - Tulip Medical Products shall:

      - Ensure appropriate methods/procedures for validation of the data provided shall be implemented;

      - Periodically verify the correctness of all of the data relevant to devices they have placed on the market, except for devices that are no longer available on the market;

      - Update the relevant UDI database record within 30 days of a change being made to an element, which does not require a new UDI-DI

Note: Data relating to devices that are no longer available on the market shall be retained in the UDI database.

  - Rules for Specific Device Types (as applicable)

    - Custom-made devices (MDR only)

      - The need for UDI does not apply to custom-made devices per MDR Article 27

    - Implantable Devices

      - Must be identified, or marked using AIDC, with a UDI (UDI-DI + UDI-PI), at their lowest level of packaging (‘unit packs’).

      - The UDI-PI must have at least the following:

        - the serial number for active implantable devices,

        - the serial number or lot number for other implantable devices.

      - The UDI of the implantable device shall be identifiable prior to implantation.

    - Reusable Devices Requiring Cleaning, Disinfection, Sterilization or Refurbishing Between Uses

      - The UDI shall be placed on the device and be readable after each procedure to make the device ready for the next use.

      - The UDI-PI characteristics, such as the lot or serial number shall be defined by the manufacturer.

    - Systems and Procedure Packs

      - The system or procedure pack must be identified with a UDI, including both UDI-DI and UDI-PI.

      - Device contents shall bear a UDI carrier on their packaging or on the device itself. Exemptions:

- Individual single-use disposable devices, the uses of which are generally known to the persons by whom they are intended to be used, which are contained within a system or procedure pack, and which are not intended for individual use outside the context of the system or procedure pack, shall not be required to bear their own UDI carrier;

- Devices that are exempted from bearing a UDI carrier on the relevant level of packaging shall not be required to bear a UDI carrier when included within a system or procedure pack.

      - Placement of the UDI carrier on systems or procedure packs shall be:

- affixed to the outside of the packaging,

- readable, or, in the case of AIDC, scannable, whether placed on the outside of the packaging of the system or procedure pack or inside transparent packaging.

    - Configurable Devices

      - A UDI shall be assigned to the configurable device in its entirety and shall be called the configurable device UDI.

        - The configurable device UDI-DI shall be assigned to groups of configurations, not per configuration within the group.

Note: A group of configurations is defined as the collection of possible configurations for a given device as described in the TD.

        - A configurable device UDI-PI shall be assigned to each individual configurable device.

      - The carrier of the configurable device UDI shall be placed on the assembly that is most unlikely to be exchanged during the lifetime of the system and shall be identified as the configurable device UDI.

      - Each component that is considered a device and is commercially available on its own shall be assigned a separate UDI.

    - Software Device

      - The UDI shall be assigned at the system level of the software.

a) 	Only software which is commercially available on its own and software which constitutes a device in itself shall be subject to that requirement.

      - The software identification shall be considered to be the manufacturing control mechanism and shall be displayed in the UDI-PI.

      - A new UDI-DI shall be required whenever there is a modification that changes:

- the original performance;

- the safety or the intended use of the software;

- interpretation of data.

      - Such modifications include new or modified algorithms, database structures, operating platform, architecture or new user interfaces or new channels for interoperability.

      - Minor software revisions shall require a new UDI-PI and not a new UDI-DI.

        - Minor software revisions are generally associated with bug fixes, usability enhancements that are not for safety purposes, security patches or operating efficiency.

        - Minor software revisions shall be identified by a manufacturer-specific form of identification.

      - UDI placement criteria for software:

- where the software is delivered on a physical medium, e.g. CD or DVD, each packaging level shall bear the human readable and AIDC representation of the complete UDI. The UDI that is applied to the physical medium containing the software and its packaging shall be identical to the UDI assigned to the system level software;

- the UDI shall be provided on a readily accessible screen for the user in an easily-readable plain-text format, such as an ‘about’ file, or included on the start-up screen;

- software lacking a user interface such as middleware for image conversion, shall be capable of transmitting the UDI through an application programming interface (API);

- only the human readable portion of the UDI shall be required in electronic displays of the software. The marking of UDI using AIDC shall not be required in the electronic displays, such as ‘about’ menu, splash screen etc.;

- the human readable format of the UDI for the software shall include the Application Identifiers (AI) for the standard used by the issuing entities, so as to assist the user in identifying the UDI and determining which standard is being used to create the UDI.

  - Determine the accredited ‘Issuing Agency / Entity’ to be used

    - 6.6.1. The Commission (per MDR Art. 120(12) and Art. 27(2)) has designated the following issuing entities that operate an electronic registration system for assignment of company identification numbers and device identifiers:

- GS1 US, Lawrenceville, NJ www.gs1.org

- Health Industry Business Communications Council (HIBCC), Phoenix, AZ www.hibcc.org

- ICCBBA, San Bernardino, CA www.iccbba.org

    - After the application and fee is made to the Issuing Entity, Tulip Medical Products is assigned a unique ‘company identification number’ from that is used to identify Tulip Medical Products and this number becomes part of the DI number.

  - Technical Documentation (TD)

    - The TD will contain the UDI information including the UDI-DI, refer to QSP 9.1-3.

  - Changes to Devices

    - A new UDI-DI shall be required whenever there is a change that could lead to misidentification of the device and/or ambiguity in its traceability; in particular, any change of one of the following UDI database data elements shall require a new UDI-DI:

- name or trade name,

- device version or model,

- labelled as single use,

- quantity of devices provided in a package,

- critical warnings or contra-indications; e.g. containing latex or DEHP.

    - New finished device packaging requires a new DI number to be assigned.  An example is device packaging that now includes accessories or changes to the packaged quantities.

    - Information that is incorrect or potentially misleading must be updated immediately in Eudamed.

    - Tulip Medical Products will update the UDI database record within 30 days of a change being made to an element, which does not require a new UDI-DI.

  - Discontinued Devices

    - When finished devices are discontinued, the UDI number must not be used again for a different model or version of a device.

    - Enter the discontinue date on the EU UDI Number Listing, Form 9.1-6-1 with the status that the product is no longer in commercial distribution.

    - Ensure that the UDI number from a discontinued product is retired from use; indicate on Form 9.1-6-1that the UDI number is no longer active.

- QUALITY RECORDS

7.1 	All Unique Device Identification (UDI) documentation will be maintained by the PRRC or designee in accordance with QSP 4.2-1 in accordance with established requirements.

7.2 	Retain the information contained in the EU UDI Number Listing, Form 9.1-6-1, to support that the unique numbers have been identified to each of the finished devices version or model numbers.

7.3 	The information contained in the UDI central file must be kept current and updated as necessary to match the product information contained in the TD.

7.4 	Current labeling showing the UDI information shall be maintained as part of the TD.

END OF DOCUMENT