TECHNICAL DOCUMENTATION - MDR

Approvals

V.P Regulatory Affairs 						             Date

Regulatory Affairs Coordinator					             Date

Document Control Administrator					              Date

| Rev | Date | Change Description |
| --- | --- | --- |
| 0 |  | Initial release. Implementation of document for compliance to MDD 93/42/EEC. |
| 1 |  | Updated references to ISO 13485:2012 |
| 2 |  | Updated sec. 2, “REFERENCE DOCUMENTS” section to reflect current publication dates. |
| 3 | 12/06/2018 | Complete rewrite and reformat. |
| 4 | 11/26/2019 | Added signature block & title. Reformatted entire document font. Revised Attachment A – Declaration of Conformity |
| 5 | 01/14/2021 | Updated attachment A: 9.1 Form 7.5-1-7-4 Rev 3 Declaration of Conformity Class II |
| 6 | 2/29/2021 | Update Responsibility for Biological Review requirement – Section 6.3 - TF finding 3.7.1.01 Update requirement to add Management Review reference for annual risk review to TD – Section 7.4.3.4.5.10 – TF finding 3.6.01 Update requirement for device classification – Section 7.4.2.5 – TF finding 3.3.01 Update requirement to add “off the shelf” accessories and their information to the TD – Section 7.4.3.5.2.1 – TF finding 3.1.01 |
| 7 | 10/13/21 | Add clarification to TF that IFU translations to their perspective countries are the responsibilities of the distributors. Section 7.10.2 |
| 8 | 5/23/25 | Update in compliance with MDR 2017/745 Regulatory Framework Migration: Complete transition from compliance with MDD 93/42/EEC to MDR 2017/745, including restructuring of content around MDR Annexes I, II, III, IV, V, VI, VIII, XIII, and XIV. Scope Clarification: Clearly defined applicability of the procedure to Class IIa, non-active, non-measuring, non-electronic, single-use medical devices, with conformity assessment conducted under Annex IX or XI with Notified Body involvement. Reference Document Section Expanded: Updated and expanded the list of reference documents to include MDR annexes, ISO 14971:2019, ISO 15223-1, EN 1041, MedTech Europe symbol guidance, and internal QSPs/forms relevant to TD, CER, PMS, and UDI. MDCG Guidance Added: Inserted citations to MDCG 2020-5, 2020-6, and 2020-13 for PMCF and Clinical Evaluation alignment with MDR expectations. Labeling and IFU Language Updated: Clarified that translation responsibilities lie with distributors and that translation quality must preserve the meaning of medical instructions. CE mark sizing guidance from Annex V integrated. Terminology and Definitions Aligned with MDR Article 2: Overhauled and expanded the definitions section to align directly with MDR terminology. GSPR Section Updated: Full integration of Form 9.1-2-3 GSPR Checklist, including justification requirements, document cross-references, and harmonized standards application. Record Retention Reworded: Updated the language to match MDR Article 10(8): "TD must be retained for the lifetime of the device plus 10 years after the last product has been placed on the market." Non-Applicable Content Removed or Marked (Software/Electronics): Removed references to software validation (EN 62304) and electrical testing (EN 60601), which were not applicable to Tulip’s product scope. Appendix A Reformatted: Provided a clear MDR Annex II-based recommended table of contents for the Technical Documentation File, tailored to Class IIa devices. "NOTE:" Sections Audited and Updated: Reviewed all “NOTE:” statements to remove outdated language, clarify record retention responsibilities, and affirm device-specific exclusions (e.g., software, measuring functions). |
| 9 | 7/14/2025 | Updated Section 5 (GSPR) to clarify the required assessment of Common Specifications (CS) in accordance with MDR Articles 2(71) and 9. Added language referencing the revised GSPR Checklist (Form 9.1-2-3 Rev 1), which now includes a column for documenting applicable CS or other regulatory references. Updated Section 6 (Risk Management) to clarify RA responsibilities for annual surveillance of CS, per QSP 7.3-3 Section 7.8. Added cross-reference to Form 7.3-3-4 (Regulatory Update Log) and Management Review (QSP 5.6-1) for regulatory oversight. |

- PURPOSE

The purpose of this procedure is to detail the mechanism used by Tulip Medical Products to document and control the development, maintenance and monitoring of the Technical Documentation (TD) that is required for medical device products sold within the European Community.  This procedure details the requirements for assembling the TD for the Medical Devices Regulation (MDR) (EU) 2017/745.

Note: that per Art 10(4), Manufacturers of custom-made devices are exempt from establishing TD for those devices. Manufacturers of custom-made devices shall draw up, keep up to date and keep available for competent authorities’ documentation in accordance with Section 2 of Annex XIII.

- SCOPE

This QSP applies to Class IIa medical devices that are non-active, non-measuring, non-electronic, and software-free. Devices include sterile, single-use cannulas and accessories. Conformity assessment is conducted under MDR Annex IX (or XI, as applicable) with oversight from the Notified Body.

- Reference DOcuments

| Medical Device Regulation (MDR) (EU) 2017/745 Annexes I, II, III, IV, V, VI, VIII, XIII, XIV |  |
| --- | --- |
| ISO 14971:2019 Medical devices – Application of risk management to medical devices |  |
| EN ISO 15223-1:2016 Medical devices – Symbols to be used with medical device labels, labeling and information to be supplied – Part 1: General requirements |  |
| EN 1041:2008 Information Supplied by the Manufacturer with Medical Devices |  |
| MedTech Europe - Use of Symbols to Indicate Compliance with the MDR (May 2019) www.medtecheurope.org/resource-library/use-of-symbols-to-indicate-compliance-with-the-mdr/ |  |
| Form 7.5-1-7-3 | Declaration of Conformity – Class I |
| Form 7.5-1-7-4 | Declaration of Conformity – Class II |
| Form 9.1-2-3 | General Safety and Performance Requirements (GSPR) Checklist (Annex I) |
| Form 7.3-3-4 | Regulatory Update Log |
| QSP 4.2-1 | Document Control |
| QSP 7.3-1 | Design and Development |
| QSP 7.3-3 | Risk Management |
| QSP 9.1-1-1 | MDR Compliance Manual |
| QSP 9.1-4 | Development of Clinical Evaluation Report |
| QSP 9.1-7 | Post Market Surveillance (PMS) and Post-Market Clinical Follow-Up (PMCF) |
| QSP 9.1-6 | Medical Device Reporting - MDR |
| QSP 9.1-6-1 | Unique Device Identification (UDI) - MDR |
| Appendix A | Recommended Format for the TD Table of Contents |

- Definitions

Accessory for a medical device (Art. 2(2)): an article which, while not being a medical device in its own right, is intended by the manufacturer to be used together with one or several particular medical device(s) to specifically enable the medical device(s) to be used in accordance with its/their intended purpose(s) or to specifically and directly assist the medical functionality of the medical device(s) in terms of its/their intended purpose(s).

Authorized representative (Art. 2(32)): any natural or legal person established within the Union who has received and accepted a written mandate from a manufacturer, located outside the Union, to act on the manufacturer's behalf in relation to specified tasks with regard to the latter's obligations under the MDR.

Benefit-risk determination (Art. 2(24)): the analysis of all assessments of benefit and risk of possible relevance for the use of the device for the intended purpose, when used in accordance with the intended purpose given by the manufacturer.

CE marking: a marking by which a manufacturer indicates that a device is in conformity with the applicable requirements set out in the MDR and other applicable EU legislation providing for its affixing.

Clinical evaluation (Art. 2(44)): a systematic and planned process to continuously generate, collect, analyze and assess the clinical data pertaining to a device in order to verify the safety and performance, including clinical benefits, of the device when used as intended by the manufacturer.

Common specifications: a set of technical and/or clinical requirements, other than a standard, that provides a means of complying with the legal obligations applicable to a device, process or system.

Conformity assessment (Art. 2(40)): the process demonstrating whether the requirements of this Regulation relating to a device have been fulfilled.

Economic Operator (EO): a manufacturer, an authorized representative, an importer, a distributor or the person referred to in Art. 22(1) and 22(3).

Eudamed: the European database on medical devices that integrates different electronic systems to collate and process information regarding devices on the market and the relevant economic operators, certain aspects of conformity assessment, notified bodies, certificates, clinical investigations, vigilance and market surveillance. {Currently delayed until May 2022}

Harmonized standard: a European standard as defined in point (1)(c) of Art. 2 of Regulation (EU) No 1025/2012.

NOTE:	 There are currently no Harmonized Standards for the MDR; in this absence, current Harmonized Standards under the MDD or more recent international versions can likely be used until standards become harmonized under the MDR. Refer to QSP 4.2-1 for monitoring of MDR-related guidance documents and standards.

Intended purpose (Art. 2(12)): the use for which a device is intended according to the data supplied by the manufacturer on the label, in the instructions for use or in promotional or sales materials or statements and as specified by the manufacturer in the clinical evaluation.

Instructions for use (Art. 2(14)): the information provided by the manufacturer to inform the user of a device's intended purpose and proper use and of any precautions to be taken.

Invasive device (Art. 2(6)): any device which, in whole or in part, penetrates inside the body, either through a body orifice or through the surface of the body.

Label (Art. 2(13)): the written, printed or graphic information appearing either on the device itself, or on the packaging of each unit or on the packaging of multiple devices.

Making available on the market (Art. 2(27)): any supply of a device, other than an investigational device, for distribution, consumption or use on the Union market in the course of a commercial activity, whether in return for payment or free of charge.

Manufacturer (Art. 2(30)): a natural or legal person who manufactures or fully refurbishes a device or has a device designed, manufactured or fully refurbished, and markets that device under its name or trademark.

Medical device (Art. 2(1): any instrument, apparatus, appliance, software, implant, reagent, material or other article intended by the manufacturer to be used, alone or in combination, for human beings for one or more of the following specific medical purposes:

- diagnosis, prevention, monitoring, prediction, prognosis, treatment or alleviation of disease,

- diagnosis, monitoring, treatment, alleviation of, or compensation for, an injury or disability,

- investigation, replacement or modification of the anatomy or of a physiological or pathological process or state,

- providing information by means of in vitro examination of specimens derived from the human body, including organ, blood and tissue donations,

and which does not achieve its principal intended action by pharmacological, immunological or metabolic means, in or on the human body, but which may be assisted in its function by such means.

The following products shall also be deemed to be medical devices:

- device for the control or support of conception.

- products specifically intended for the cleaning, disinfection or sterilization of devices as referred to in MDR Article 1(4) and of those referred to in the first paragraph of this point.

Performance (Art. 2(22)): ability of a device to achieve its intended purpose as stated by the manufacturer.

Placing on the market (Art. 2(28)): the first making available of a device, other than an investigational device, on the Union market.

Post market surveillance (PMS) (Art. 2(60)): all activities carried out by manufacturers in cooperation with other EOs to institute and keep up to date a systematic procedure to proactively collect and review experience gained from devices they place on the market, make available on the market or put into service for the purpose of identifying any need to immediately apply any necessary corrective or preventive actions.

Procedure pack (Art. 2(10)): a combination of products packaged together and placed on the market with the purpose of being used for a specific medical purpose.

Putting into service (Art. 2(29)): the stage at which a device, other than an investigational device, has been made available to the final user as being ready for use on the Union market for the first time for its intended purpose.

Risk (Art. 2(23)): the combination of the probability of occurrence of harm and the severity of that harm.

Single-use device (Art. 2(8)): a device that is intended to be used on one individual during a single procedure.

System (Art. 2(11)): a combination of products, either packaged together or not, which are intended to be inter-connected or combined to achieve a specific medical purpose.

Technical Documentation (TD): the name given to the EU regulatory documentation / file for all medical devices.

Unique device identifier (UDI) (Art. 2(15)): a series of numeric or alphanumeric characters that is created through internationally accepted device identification and coding standards and that allows unambiguous identification of specific devices on the market.

- ABBREVIATIONS

AR – Authorized Representative

CA – Competent Authority

CEP – Clinical Evaluation Plan

CER – Clinical Evaluation Report

CS – Common specification

DoC – Declaration of Conformity

EO – Economic Operator

GSPR – General Safety and Performance Requirements

IFU – Instructions for Use

MDR – Medical Devices Regulation

PMCF – Post Market Clinical Follow-up

PMS – Post Market Surveillance

PRRC - Person Responsible for Regulatory Compliance

QMS – Quality Management System

TD – Technical Documentation

- Responsibilities

  - The PRRC is responsible for ensuring that

    - The conformity of the device is appropriately checked, in accordance with the QMS under which the devices are manufactured, before a device is released.

    - The td and EU doc are drawn up and kept up to date; and

    - The PMS obligations are complied with in accordance with MDR Art. 10(10).

  - The product development team and quality will assist with the assembly of the td during the design and development of products.

  - It is the responsibility of the product development team for design and development to provide the necessary information and safety analysis for incorporation into the td.

- Procedures

  - TD is only required for medical devices that are distributed in the European Economic Area (EEA) and Switzerland.

  - Medical Device Designation and Device Classification

    - Each product is assessed to determine whether it meets the definition of a medical device or accessory as defined in Art. 2.

    - [Manufacturer’s] medical device(s) has/have been classified as Class I self-certified per MDR Annex VIII Classification Rules.

    - If changes are made to devices or new devices developed, review for the appropriate medical device classification, per MDR Annex VIII Classification Rules. If assistance is needed:

- Consult existing guidance,

- Consult with selected CA for a ruling, and/or

- Consult with our AR for input.

  - Conformity Assessment Procedures (Art. 52)

    - Class Class IIa devices require TD in accordance with Annexes II and III and are subject to conformity assessment under Annex IX (or XI, if applicable) with involvement of a Notified Body.

    - The recommended format for the TD File is described below, identifying the different sections to be included. The table of contents for the recommended format is included as Appendix A.

    - The format of the TD File does not have to be exact, although the information must be presented in a clear, organized, readily searchable and unambiguous manner and shall include in particular the elements listed in Annex II.

    - Cover Sheet contents

      - The Cover Page will indicate the document is the TD File, medical device product line (family or product), revision level, issue date, and company name.

      - The revision page identifies the revision level of the file as well as provides a short description for any subsequent revisions.

  - SECTION 1 - ADMINISTRATIVE SECTION

    - Name and address of the manufacturer.

    - Name and address of the AR, if relevant (manufacturer is outside of the EU).

    - Declaration of Conformity (DoC) (Annex IV).

      - There is no required format for the DoC as long as all of the required information is present

      - A template DoC is available: Form 7.5-1-7-3 (Class I self-certified).

      - The DoC must contain a statement that the device is in conformity with the MDR. If other EU Directives / Regulations are also applicable, these must also be referenced on the same DoC.

      - A copy of the DoC will be provided to the AR for the product as well as each Importer / Distributor of the product in the EU.

      - To reduce possible customs delays, a copy of the DoC can be provided with each shipment – this is not mandatory.

    - The status of the QMS is discussed as well as key QMS procedures. A list of procedures can be provided as an appendix to the TD File.

    - A statement regarding TD File development, review and changes should be included.

  - SECTION 2 – DEVICE DESCRIPTION AND SPECIFICATION, INCLUDING VARIANTS AND ACCESSORIES

    - Device Description

      - Trade name and general description of the device / device family including its intended purpose and intended users.

        - To be considered as a product family, all the products considered for one TD File must share design, function, and components. If the TD is to cover a family of related devices, then the justification must be included why these devices are a family.

      - Provide the Basic UDI-DI as referred to in Annex VI Part C assigned by the manufacturer to the device, as soon as identification of this device becomes available based on a UDI system, or otherwise a clear identification by means of product code, catalogue number or other unambiguous reference allowing traceability.

      - Include or refer to an up to date list of all UDIs that have been assigned to the devices in the device family.

      - Describe the intended patient population and medical conditions, indications for use, contraindications, and warnings.

- These must match that indicated in the instructions for use.

      - Principles of operation and mode of action scientifically demonstrated if necessary.

      - A rationale for the qualification of the product as a medical device.

      - The risk classification of the device and the justification for the classification rule(s) applied.

      - Explanation of novel features.

      - List of accessories, other medical devices, other products not medical devices which are intended to be used with the device. These can be provided a list or in a table format. Identify all accessories that are used for the particular device or family of devices.

      - List of various configurations / variants / models of the device.

      - Description of key functional elements.

- Describe the device(s)’ parts/components, formulation, composition, functionality and, where relevant, the qualitative and quantitative composition.

- Where appropriate, this can include labeled pictorial representations (e.g. diagrams, photographs, and drawings), clearly indicating key parts/components, including sufficient explanation to understand the drawings and diagrams.

      - Description of raw materials incorporated into key functional elements which have either direct or indirect contact with the human body.

      - List technical specifications, such as features, dimensions and performance attributes, of the device and any variants/configurations and accessories that would typically appear in the product specification made available to the user, for example in brochures, catalogues and similar publications.

    - Reference to Similar and Previous Generations of the Device

      - Review manufacturer’s previous generations of the device, if existent.

      - Discussion of similar devices available on local or international markets.

  - SECTION 3 – INFORMATION SUPPLIED BY THE MANUFACTURER

    - Include complete set of labeling: labels on the device and packaging, Instructions for Use (IFU), and promotional materials.

    - The Instructions for Use and labels are be developed in English, reviewed by the AR for completeness, and then translated to the language(s) where the device will be placed on the market.

    - Consider the use of internationally accepted symbols from EN ISO 15223-1 and the MedTech Europe guidance in lieu of text.  Refer to EN 1041 for additional information on the Instructions for Use.

NOTE: 	Refer to QSP 9.1-1-1 MDR Compliance Manual for further details on labeling.

    - As translation of the product labeling can create errors, be careful to use recognized and medical device experienced translation services to ensure the meaning of the instructions are not inadvertently modified.

    - Refer to Annex V of the MDR for regulations on the size proportions of the CE mark.

  - SECTION 4 – DESIGN AND MANUFACTURING INFORMATION

    - Device Design

      - General description of design stages, with flow chart, as applicable. These can be provided as a list or in a table format.

      - Identify those design specifications that are applicable to the development of this medical device.

NOTE:	The proof of burden is on the manufacturer.  If required for any litigation, evidence must be shown that risks have been adequately assessed and test data shown to provide confirmation of design.

    - Manufacturing Process

      - General description of manufacturing processes. Process flow chart showing overview production, assembly, continuous monitoring, final product testing, packaging, finished medical device or descriptions of each.

      - Provide validation information for the manufacturing processes, as applicable.

    - Design and Manufacturing Sites

      - Identify all sites, including suppliers and sub-contractors, where design and manufacturing activities are performed.

- If these sites possess QMS or similar certifications / registrations, this will be indicated and the certificates referenced.

NOTE:	 These should be reviewed as part of the annual TD review to ensure they are still valid.

  - SECTION 5 – GENERAL SAFETY AND PERFORMANCE REQUIREMENTS (GSPR)

    - The GSPR are taken from MDR Annex I. Refer to Form 9.1-2-3 GSPR Checklist.

    - Common Specifications (CS) as defined in MDR Article 2(71) and Article 9 shall be assessed during the preparation and review of the GSPR checklist.
Where applicable, CS must be used in place of or in addition to harmonized standards. The GSPR checklist (Form 9.1-2-3) includes a dedicated column for documenting relevant CS, MDCG guidance, or other authoritative references.
Responsibility for identifying applicable CS rests with the Regulatory Affairs team in consultation with Risk Management (QSP 7.3-3), and justification for inclusion or exclusion shall be documented in the checklist.

    - Each element of the GSPR is assessed for the product / family of products. In each requirement, the intention is to identify how the manufacturer complies with each relevant requirement as well as reduce risk as reduce risks as far as possible without adversely affecting the benefit-risk ratio.

    - Obtain Form 9.1-2-3 GSPR Checklist and complete the data for each item:

      - Compliance Column:

      - Each GSPR must be assessed to ascertain if it is applicable to the subject device/device family and if the device is in compliance for each requirement (Y/N/NA). The completed GSPR must be fully compliant.

      - Applied Standards/CSs/Guidance Documents Column:

- In this column list any/all standards, CSs, guidance documents, etc., that device is compliant with.

      - Complying Documents Column:

- List the precise identity of all specific documents / test reports / etc. (including version, at least with the first time listed) that demonstrate compliance with each applicable GSPR as well as the harmonized / other standards / CS / guidance documents referenced.

      - Location Column:

- Identify the location of the documents, test reports, etc. listed in the previous column.

    - Each GSPR must be justified.  If a requirement is not applicable to your device, then state NA and why it is not applicable in the ‘Applied Standards / CSs / Guidance Documents’ column.  It is not suitable to make a claim of non-applicability without providing a justification.

    - Assign the GSPR a document number and revision level according to Document Control process or process for numbering product related documents.

  - SECTION 6 – BENEFIT-RISK ANALYSIS AND RISK MANAGEMENT

    - Include a summary of risks identified during risk analysis process and discussion how risks have been controlled or mitigated to an acceptable level. Conducting a risk analysis and risk management is mandatory.  EN ISO 14971 is the applicable standard to follow.  Refer to QSP 7.3-3 Risk Management.

    - Regulatory Affairs shall perform annual surveillance to identify changes in Common Specifications and related regulatory guidance, as described in QSP 7.3-3 Section 7.8. Any impact to GSPR conformity or residual risk shall be assessed collaboratively with the Risk Management team and documented in Form 7.3-3-4 – Regulatory Update Log.
The Technical Documentation shall be updated accordingly, and Management Review (QSP 5.6-1) shall be used to ensure oversight and implementation.

    - Briefly describe the risk management process in this section and reference the location of the Risk Management File.

    - The Risk Analysis must be conducted by personnel with knowledge and experience appropriate to the assigned task.  This includes as appropriate, experience and knowledge of the medical device and its use and risk management CER/PMF techniques.  Records of the appropriate qualifications must be maintained within the QMS.

  - SECTION 7 – PRODUCT VERIFICATION AND VALIDATION

    - Pre-Clinical Data

      - Summary of results of verification and validation tests and/or studies taken to demonstrate conformity of the device with the requirements of the MDR and in particular the GSPR.

      - If applicable, information on the device related to the following will be summarized (as applicable):

        - Biocompatibility

- List materials in direct or indirect contact with the patient or user

- Summary of biocompatibility testing to characterize the physical, chemical, toxicological and biological response of materials in direct or indirect contact with the patient or user (testing is to be performed on samples from finished and sterilized device, if applicable).

- Refer to EN ISO 10993-1 [and other applicable standards based on testing performed].

        - Physical, Chemical and Microbiological Characterization.

- Summary of testing to characterize the physical, chemical, and/or microbiological characteristics of relevant device materials (testing is to be performed on samples from finished and sterilized device, if applicable).

        - Packaging Qualification

Design and development requirements for packaging systems:

- General aspects

- Design

- Packaging-system performance testing

- Stability testing

These aspects will be supported by either test data or a rationale supported by historic data. Clear drawings of the total packaging system configuration will be added or referenced.

The validation will be documented in a report, summarizing the results of all scenarios tested.

        - Shelf Life

The following aspects need to be addressed, as applicable:

- Expiry period (shelf life);

- Shelf life protocol.

- Shelf life report; and

- Special storage requirements (if applicable).

The stability of the product and packaging system shall be demonstrated for the shelf life (expiration date) claimed by the manufacturer. This can be done by testing after real-time or accelerated aging. Alternatively, a rationale can be given supported by historic data.

        - Lifetime of Device

Describe lifetime of product with rationale to support (e.g. product life-cycle testing, analysis of components or device, history on similar devices, etc.).

The MDR record retention requirement for TD is the actual lifetime of the device plus 10 years for non-implantable devices, after the last date of manufacture.

        - Performance and Safety Testing

Describe (in text or table) or reference performance and safety testing performed, not previously described above. Usability Studies may be one type of testing performed for this category.

      - Identify what harmonized standards and common specifications were used to verify that the product achieves compliance with the EU requirements. It is helpful to determine specifications that are appropriate for a particular medical device;

http://ec.europa.eu/growth/single-market/european-standards/harmonised-standards/

        - Although harmonized standards are voluntary, if applied they provide the “presumption of conformity” to the GSPR. We therefore recommend that European standards are reviewed and, where applicable, applied.

      - If the test data is held in the Design History File or other file or record, then state this in this section.

      - Where no new testing has been undertaken, document a rationale for that decision. An example of such a rationale would be that biocompatibility testing on identical materials was conducted when those materials were incorporated in a previous version of the device that was sold in the EU.

    - Clinical Data

      - The manufacturer shall specify and justify the level of clinical evidence necessary to demonstrate conformity with the relevant GSPR. That level of clinical evidence shall be appropriate in view of the characteristics of the device and its intended purpose.

Therefore, where the demonstration of conformity with GSPR based on clinical data is not deemed appropriate, provide adequate justification based on the results of the risk management and on consideration of the specifics of the interaction between the device and the human body, the clinical performance intended, and the claims. In such a case, substantiate in the TD because adequate demonstration of conformity with GSPR can be based on the results of non-clinical testing methods alone, including performance evaluation, bench testing and pre-clinical evaluation.

      - Describe (i.e. reference and/or summarize) the clinical tests and/or safety and efficacy testing that have been performed to validate the medical device. This section needs to include data to support the safety and performance of the device. If clinical tests have been performed, summary can be added here and/or reference location of clinical data. Refer to QSP 9.1-4 Development of Clinical Evaluation Report.

NOTE: 	Include reference to the completed CEP and CER for the device.

NOTE: 	Refer to Annex XIV Part A for further details on clinical evaluation.

    - Post Market Clinical Follow-up (PMCF)

      - Provide summary of PMCF Plan and PMCF Report (MDCG 2020-6) referred to in Annex XIV Part B or a justification why a PMCF is not applicable.

      - Make reference to the PMCF Plan and PMCF Report, as applicable.

      - Refer to QSP 9.1-7 PMS & PMCF.

Note: PMS data is to be provided in Section 8 below.

    - Additional Information Required in Specific Cases

NOTE:	Sections 7.10.4.1–7.10.4.6 may not apply to Tulip’s current product scope; where not applicable, this will be documented and justified in the TD.

      - Medicinal Substances

      - Human and Animal Tissue or Cells

      - Substances Intended for Absorption or Local Disbursement into Human Body

      - Devices Containing CMR or Endocrine-Disrupting Substances

      - Sterile Devices or Defined Microbiological Condition

      - Devices with a Measuring Function

      - Devices Intended to be Connected to Other Devices

        - Description of this combination/configuration including proof that it conforms to the GSPR when connected to any such device(s) having regard to the characteristics specified.

    - Post Market Surveillance

      - Provide summary of the PMS Plan and PMS Report referred to in Annex III, Art 83, 84 & 85.

      - Make reference to the PMS Plan and PMS Report.

      - Refer to QSP 9.1-7 PMS & PMCF, used to create the PMS Plan and Report.

    - Include applicable Appendices, as referenced in the various sections above.

  - QUALITY RECORDS

    - The TD is maintained in a central location and identified by the control number assigned.

    - Each of the sections of the TD must be complete or reference the location where the other documents are located.

    - During the annual review of the TD, conduct a search to determine if any new standards, guidance documents, or CSs are applicable and to verify that the most current revisions are listed in the TD File and are used as necessary. Also refer to QSP 4.2-1.

    - The TD for Class IIa devices must be retained for the lifetime of the device plus 10 years after the last product has been placed on the market, per MDR Article 10(8).

NOTE: 	Must factor in the device life to retention timeframe.

Appendix A:  Recommended* Format for the TD – Table of Contents

| SECTION | DESCRIPTION |
| --- | --- |
| Cover Page | Cover page for particular product family |
| Revision Page | Identifies the revision level of each section / reason for revision |
| 1 | Administrative Information - Name and address of the Manufacturer - Name and address of the AR, if relevant - DoC / Conformity Assessment Route - Quality System & Regulatory Requirements - TD File Development & Changes |
| 2 | Device Description and Product Specification, Including Variants and Accessories |
| 2.1 | Device Description and Specification - Device Trade Name and Description - Basic Unique Device Identifier - All UDIs that have been assigned (list or refer to document) - Intended Use - Patient Population and Medical Conditions - Indications - Contraindications - Warnings - Principles of Operation - Medical Device Rationale - Risk Classification of the Device and Justification for the Classification - Novel Features - Device Accessories - Device Configurations/Variants - Key Functional Elements - Raw Materials - Technical Specifications |
| 2.2 | Previous and Similar Generations of the Device - Overview of Previous Generation or Generations of the Device - Overview of Identified Similar Devices |
| 3 | Information Supplied by the Manufacturer - Device Label - Device IFU - Device CE Marking |
| 4 | Design and Manufacturing Information - Device Design - Device Manufacturing - Identification of Sites |
| 5 | GSPR |
| 6 | Benefit/Risk Analysis and Risk Management - Benefit/Risk Analysis - Risk Management |
| 7 | Product Verification and Validation |
| 7.1 | Pre-Clinical and Clinical Data - Pre-Clinical Data - Biocompatibility Testing - Physical, Chemical, and Microbiological Characterization - Tulip Medical Products currently manufactures no devices with software or programmable components. - Device Package & Shelf-Life Validation/Lifetime of Device - Device Performance Testing - Applied Standards Clinical Data PMCF |
| 7.2 | Additional Information Required in Specified Cases - Medicinal Substances - Human and Animal Tissue or Cells - Substances Intended for Absorption or Local Dispersement into Human Body - Devices Containing CMR or Endocrine-Disrupting Substances - Sterile Devices or Defined Microbiological Condition. |
| 8 | Post-Market Surveillance (PMS) - PMS Plan - PMS Report |
|  | Appendices |

*This format is based on MDR, Annex II, Technical Documentation.