Tulip Medical

Quality System Procedure

QSP 7.3-3

Risk Management

Approvals

V.P Regulatory Affairs     							Date

Regulatory Affairs Coordinator 				                                    Date

Document Control Administrator					                  Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 05/07/2013 | Initial release. |
| 1 | 06/27/2014 | Incorporate the requirement of ISO 14971:2012. |
| 2 | 06/30/2014 | Added section on Risk Reduction (section 6.4.3). |
| 3 | 09/26/2018 | Reformatted and included new signature blocks. |
| 4 | 12/12/2018 | Restored Appendix 1 and Appendix 2. Appendix 1, “Example of Ranking System for FMEA; Appendix 2, “Sample Outline of Risk Management Plan.” |
| 5 | 3/29/2021 | Add to Section 7.2.3 the following : Add Clearly defined allocation of responsibilities |
| 6 | 12/01/22 | Add to Section 7.1.5 The FMEAs are being reviewed on an ongoing basis through the non-conformance, complaints, and CAPA process |
| 7 | 3/31/25 | -3.0 Added Form 7.3-3-3 -7.7 Added Standalone Risk Management Report Requirement |
| 8 | 5/23/2025 | Updated QSP 7.3-2 to align with Regulation (EU) 2017/745 (MDR) and ISO 13485:2016. Reworded the purpose to reflect MDR language. Updated the scope to clarify the DMR as an MDR requirement. Added references to MDR and ISO 13485:2016 in the Reference Documents section. Removed all legacy references to MDD. Updated terminology to align with MDR Annex II and III regarding technical documentation. |
| 9 | 7/14/2025 | Added Section 7.9 – Regulatory Surveillance and Impact Assessment to address MDR Annex I and ISO 14971 requirements for monitoring changes to Common Specifications (CS), harmonized standards, and regulatory guidance. Integrated responsibilities for regulatory update tracking, risk impact evaluation, and management review reporting. Section 7.3.7 and Section 7.5.2 Update to align with ISO/TR 24971 Added Form 7.3-3-4 Regulatory Update Checklist Log to section 3.0 Added ISO/TR 24971 to section 3.0 Removed MR from the abbreviation and Responsibilities Added PRRC to abbreviations and Responsibilities |

- PURPOSE

To initiate and maintain risk management activities from the outset of design projects, ensuring continuous assessment and control of risks throughout the product lifecycle. This includes the identification, analysis, control, and monitoring of hazards associated with medical devices to ensure they remain within acceptable risk levels.

Risk assessments shall be applied during the development of new products, as well as for any changes to existing products that may affect their fit, form, function, or intended purpose.

- SCOPE

This procedure covers the total product life cycle, through device requirements, production, distribution, use, maintenance and obsolescence. It is intended to serve as a systematic approach to identify possible risks of a device or system and a means to show that they have been identified, addressed and controlled.

- REFERENCE DOCUMENTS

| Regulation (EU) 2017/745 on medical devices (MDR) |  |
| --- | --- |
| Food and Drug Administration 21 CFR 820 Quality System Regulation |  |
| International Standard Medical Devices – Application of risk management to medical devices ISO 14971:2019 ISO/TR 24971 - Medical devices — Guidance on the application of ISO 14971 |  |
| SOR/98/282 Health Canada Medical Device Regulations |  |
| Form 6.2-1-2 | Training Record |
| Form 7.3-3-1 | Risk Management Plan |
| Form 7.3-3-3 | Risk Management Report and Summary |
| Form 7.3-3-2 | FMEA |
| Form 7.3.3.4 | Regulatory Update Checklist Log |
| QSP 4.2-1 | Document Control |
| QSP 5.6-1 | Management Review |
| QSP 6.2-1 | Training and Qualification |
| QSP 7.3-1 | Design and Development |
| QSP 9.1-3 | Development of Technical File |
| QSP 9.1-7 | Post Market Surveillance |

- DEFINITIONS

Harm – Physical injury or damage to the health of people, or damage to property or the environment.

Hazard – Potential source of harm.

Hazardous Situation – Circumstances in which people, property, environment are exposed to one or more hazard(s).

Intended Use / Intended Purpose – Use for which a product, process or service is intended according to the specifications, instructions, and information provided by the manufacturer.

Life-Cycle – All phases in the life of a medical device, from the initial conception to final decommissioning and disposal.

Objective Evidence – Data supporting the existence or verity of something.

Post-Production – Part of the life-cycle of the product after the design has been completed and the medical device has been manufactured.

Residual Risk – Risk remaining after risk control measures have been taken.

Risk – Combination of the probability of occurrence of harm and the severity of that harm.

Risk Analysis – Systematic use of available information to identify hazards and to estimate the risk.

Risk Assessment – Overall process comprising a risk analysis and a risk evaluation.

Risk Control – Process in which decisions are made and measures implemented by which risks are reduced to, or maintained within, specified levels.

Risk Estimation – Process used to assign values to the probability of occurrence of harm and the severity of that harm.

Risk Evaluation – Process of comparing the estimated risk against given risk criteria to determine the acceptability of the risk.

Risk Management – Systematic application of management policies, procedures, and practices to the tasks of analyzing, evaluating, controlling and monitoring risk.

Safety – Freedom from unacceptable risk.

Severity – Measure of the possible consequences of a hazard.

Verification – Confirmation, through the provision of objective evidence, that specified requirements have been fulfilled.

- ABBREVIATIONS

EVP – Executive Vice President

FDA – Food and Drug Administration

FMEA – Failure Modes and Effects Analysis

PRRC- Person Responsible for Regulatory Compliance

- RESPONSIBILITIES

PRRC (Person Responsible for Regulatory Compliance): A qualified individual appointed under MDR Article 15 to ensure device conformity, maintain technical documentation, oversee post-market surveillance and vigilance, and issue statements for investigational devices.

Risk management activities are to be developed through a cross-functional team with understanding of the use of the device, production of the device, across all aspect of the product realization process as well as use by customers.

- PROCEDURE

  - Overview

    - ISO 14971 encourages thinking outside the box.  The goal of risk management is to identify potential risks or any foreseeable misuse of the device, as well as known and foreseeable hazards in normal use and misuse, to take effective steps to eliminate, reduce or mitigate the risks.

    - The risk analysis can be recorded on Form 7.3-3-2 FMEA Form, or other suitable means, for each product/product family.  This may include a design FMEA on the actual device and a process FMEA for the realization of the device for routine distribution.  The purpose of this form is to provide identification and traceability for each hazard to the risk evaluation, show proof of the implementation and verification of risk controls, and to show the assessment that the residual risks are acceptable.

    - Form 7.3-3-1 Risk Management Report and Summary can be used to describe and identify the medical device or accessory being analyzed, identify the intended use and purpose of the device, identify the person(s) who carried out the risk analysis, and document the date the analysis was performed.  This report will also summarize the overall findings of the risk analysis / FMEA.

    - The risk management documents are reviewed and updated, if applicable, at a minimum of annually. Also refer to QSP 5.6-1 Management Review and QSP 9.1-7 Post Market Surveillance.  Additional reviews and updates may be made as needed.

    - The FMEAs are being reviewed on an ongoing basis through the non-conformance, complaints, and CAPA process.

  - Risk Management Plan

    - Appendix 2 contains an outline of a Risk Management Plan that can be used.

    - Note:  Also refer to ISO 14971 Annex A – A.2.3.4 Risk management plan for additional information.

    - Prepare a risk management plan for the particular device or accessory being considered.  The plan is to include the following:

      - The scope of the plan, identifying and describing the device and the life cycle phases for which the plan is applicable

- References to the plan

- Clearly defined allocation of responsibilities

- Criteria for risk acceptability

- Requirements for review of risk management activities

- Verification activities (what type of testing)

- Approvals

- Location of risk management documents

    - See QSP 4.2-1 Document Control for minimal required approval signatures.

  - Risk Analysis

    - For the particular device or accessory, describe the intended use/purpose and any reasonably foreseeable misuse using Form 7.3-3-1 Risk Management Report and Summary or other suitable means.

    - List all qualitative and quantitative characteristics that could affect the safety of the medical device.  Annex C of ISO 14971 can be used as a guide in drawing up this list.

    - Compile a list of known or foreseeable hazards associated with the product in both normal and fault or worse-case conditions using Form 7.3-3-2 FMEA.  Identify any previously recognized hazards from similar devices, reportable events, etc.

    - The risk analysis may take the form of a design FMEA to assure that hazards with the product are identified as part of the development process.  In addition, a process FMEA identifies those risks associated with the realization of the product during routine manufacturing.  The risk analysis may include both the design FMEA and process FMEA included in one document.

    - Examples of possible hazards, foreseeable sequences of events and hazardous situations are listed in Annex E of ISO 14971.

    - For each hazard identified, estimate the risk in both normal and fault conditions using the Ranking System in Appendix 1 of this procedure.  If the probability of the occurrence of harm cannot be estimated, prepare a list of the possible consequences of the hazard.

    - Risk estimation may be conducted using either qualitative or quantitative methods, as appropriate to the device and available data. Guidance on selecting and applying these methods is provided in ISO/TR 24971:2020 Clause 5.5,

        - Risk estimation should consider the probability of occurrence of harm and the severity of potential harm. Information supporting risk estimation can be obtained from various sources, including:

- Applicable published standards and regulatory guidance

- Scientific and technical literature

- Field data, such as post-market surveillance and reported incidents from similar devices

- Usability testing involving representative users and use environments

- Clinical evidence, including clinical investigations and evaluations

- Expert judgment based on qualified experience and training

        - Personnel involved in risk estimation activities must be competent and understand their responsibilities in ensuring the accuracy, consistency, and traceability of risk evaluations, as well as their role in maintaining conformity with applicable quality procedures.

  - Risk Evaluation

    - For each identified hazard for the particular device, determine whether the estimated risk(s) is so low that risk reduction does not need to be pursued. Refer to Appendix 1 of this procedure and/or the specific product Risk Management Plan for ‘risk acceptability’.

Note:	Guidance for defining and applying these criteria is provided in ISO/TR 24971:2020, Annex C, which supports consistent evaluation of identified risks and selection of appropriate risk control measures throughout the development process.

    - If the risk is not broadly acceptable (R-1), risk reduction must be considered under the following 3 scenarios:

      - Risk is reduced to the broadly acceptable region, and has been reduced as far as possible (AFAP).

      - Risk reduction is practicable (R-2), but it cannot be reduced into the broadly acceptable region, and has been reduced as far as possible (AFAP).  Risk is considered acceptable.  Economic considerations are not acceptable to stop risk reduction.

      - Risk reduction is not reasonably practicable (R-3) – residual unacceptable risk is evaluated against the device benefits to determine its acceptability.  Device benefits should include consideration of the patient’s current state and prognosis, the likelihood of improvement or deterioration if alternative treatments are used, and the degree and likelihood of improvement with the proposed treatment.  If the benefits do not outweigh the risks, the risk is unacceptable.

Risk Acceptability

| Risk Level | RPN Range | Acceptability |
| --- | --- | --- |
| 1 | 1 through 9 | Broadly acceptable, risk has been reduced as far as possible (AFAP). |
| 2 | 10 through 27 | Risk should be reduced as far as possible (AFAP) with technical practicability and are balanced against benefit-risk. Risk is considered acceptable. Economic considerations are not acceptable to stop risk reduction. |
| 3 | 28 through 59 | Risk is unacceptable and should be reduced as far as possible (AFAP); technical practicability is balanced against benefit-risk and risk is reduced even at considerable costs. |
| 4 | > 60 | Intolerable, risk is unacceptable and must be reduced. |

  - Risk Control

    - When risk reduction is required, follow the steps listed below to control the risk(s) so that the residual risk(s) associated with each hazard may be judged acceptable.

    - Risk control measures shall be selected to reduce identified risks to an acceptable level, in accordance with the risk management policy. As recommended in ISO/TR 24971:2020 (Clause 7.1 and Annex D), the following hierarchy of risk control options shall be applied, in order of priority:

      - Inherently safe design measures (e.g., eliminating hazards or reducing their likelihood through design)

      - Protective measures incorporated into the medical device or manufacturing process (e.g., alarms, interlocks, process controls)

      - Information for safety, such as warnings, precautions, or instructions for use intended to reduce the likelihood or severity of harm

        - If, after applying available risk control measures, the residual risk is not acceptable and further reduction is not feasible, a benefit-risk analysis shall be conducted to determine whether the medical benefit outweighs the residual risk. If acceptable, the residual risk must be documented and communicated. If not, redesign or additional controls are required.

    - Implement the risk control measure(s) selected and record the measure(s) used.  Verify the implementation and effectiveness of the risk control measure(s) and record the verification.

    - Evaluate any residual risk that remains after the risk control measure(s) are applied using the ranking system below in Appendix 1 and record the results on the FMEA.

      - If the residual risk is not acceptable, apply further risk control measures as described above.

      - If the residual risk is acceptable, place all relevant information necessary to explain the residual risk(s) in the appropriate accompanying documents.

    - If the residual risk is unacceptable using the criteria established in the Risk Management Plan and further control is impractical, gather and review data and literature on the medical benefits of the intended use to determine if they outweigh the residual risk.

    - If this evidence does not support the conclusion that the medical benefits outweigh the residual risk, then the risk remains unacceptable and further corrective action must be taken.  If the medical benefits outweigh the residual risk, document the conclusion and proceed to the next step.

    - Record the relevant information necessary to explain the residual risk and the results of this evaluation.

    - Review the risk control measures to identify if new hazards have been introduced.  If any new hazards have been introduced, assess the hazard(s) by taking the steps above.  Record the results of the review on Form 7.3-3-1 Risk Management Report and Summary or use another suitable means.

    - Assure that the risk(s) from all identified hazards have been evaluated and record the results of the assessment.  Record these in Form 7.3-3-1 Risk Management Report and Summary or use another suitable means.

  - Overall Residual Risk

    - After all risk control measures have been implemented and verified, determine if the overall residual risk posed by the device is acceptable using the ranking system in Appendix 1 and/or the specific product Risk Management Plan.

    - If the overall residual risk is unacceptable using the criteria established, gather and review data and literature on the medical benefits of the intended use / intended purpose to determine if they outweigh the overall residual risk.

    - The benefit-risk statement is produced by gathering data and literature on the medical benefits of the intended use / intended purpose to determine if they outweigh the residual risk. If this evidence does not support the conclusion that the medical benefits outweigh the residual risk, then the risk remains Undesirable.  If the medical benefits outweigh the residual risk, then the risk is deemed acceptable.  Record the relevant information necessary to explain the residual risk.

    - If this evidence does not support the conclusion that the medical benefits outweigh the overall residual risk, then the risk remains unacceptable.

    - A benefit-risk statement must be compiled and finalized by the PRRC or designee for all products with hazards rated as “Undesirable”.  A written and approved decision is required to proceed on the Risk Management Report and Summary form.

    - Record the results of the overall residual risk evaluation in Form 7.3-3-3 Risk Management Report and Summary or use another suitable means.

  - Standalone Risk Management Report and Summary Requirement

    - A standalone Risk Management Report and Summary shall be created for each medical device or device family, in compliance with MDR Annex I and ISO 14971:2019, Section 4.4. This report shall:

    - Be maintained independently from the FMEA documentation.

    - Demonstrate the fulfillment of General Safety and Performance Requirements (GSPR) as per MDR Annex I.

    - This may be a single report following Form 7.3-3-3 or be broken into two parts. 1. Report and 2. Summary

    - Include:

      - Device description and intended use

      - Summary of known and foreseeable hazards

      - Risk analysis and evaluation findings

      - Implemented risk controls and verification of effectiveness

      - Benefit-risk determinations, especially where residual risks are present

      - Assessment of the overall residual risk

      - Conclusion on the acceptability of residual risks

      - Reference to relevant PMS/PMCF inputs and updates as part of risk monitoring

    - The Risk Management Report must be finalized and approved by the Management Representative and President (or designee) prior to device market release and updated post-market as part of the regular review cycle or upon significant change.

  - Risk Monitoring and Post-Production Monitoring

    - Product risk management documents are reviewed at a minimum of annually through various processes, such as internal audits, Management Reviews and feedback from complaints and post market surveillance.

    - Evaluate post market surveillance information, see QSP 9.1-7 Post Market Surveillance, for possible relevance to safety, especially the following:

      - If previous unrecognized hazards are present

      - If the estimated risk(s) arising from the hazard are no longer acceptable

      - If the original assessment is otherwise invalidated

    - If any of the above conditions exist, feedback the results of the evaluation as an input into the risk management process and determine whether a new design project should be initiated, see QSP 7.3-1 Design and Development.

    - If there is a potential that the residual risk or its acceptability has changed, evaluate the impact on previously implemented risk control measures.

    - Update the product’s Risk Management Plan, the product FMEA / risk analysis and Risk Management Report and Summary as needed.

  - Regulatory Surveillance and Impact Assessment

    - Tulip Medical shall maintain a proactive process to identify and evaluate external regulatory updates that may impact device safety, performance, or conformity.

This includes:

    - Common Specifications (CS) issued under MDR Article 9 and Annex I

    - Harmonized Standards

    - MDCG Guidance Documents

    - Updates to ISO 14971, ISO 13485, and other recognized standards

    - Relevant international regulatory guidance (e.g., FDA, Health Canada)

    - The Regulatory Affairs (RA) department shall perform surveillance on an annual basis at minimum, and upon notification of updates from regulatory bodies or standard organizations. Information sources may include:

      - European Commission Publications (NANDO, Official Journal)

      - ISO/IEC standard update notifications

      - MDCG document releases

      - FDA and Health Canada updates

      - Industry publications and Notified Body communications

    - Upon identification of a regulatory change, RA shall assess the impact to the following:

      - Device safety and performance (via risk analysis)

      - Technical Documentation and GSPR checklist (Form 9.1-2-3 Rev 0)

      - Instructions for Use (IFU), labeling, and packaging

      - Post-Market Surveillance (QSP 9.1-7) and PMCF plans

      - Clinical Evaluation Reports

      - Quality System Procedures (e.g., QSP 9.1-3)

The assessment shall be documented in the Regulatory Update Checklist Log (Form 7.3-3-4).

    - If the update introduces a new or modified requirement that may affect risk controls, clinical evidence, or conformity, RA shall initiate a Risk Management File (RMF) review. This includes:

      - Reassessment of the applicable FMEAs

      - Updates to the Standalone Risk Management Report and Summary (Form 7.3-3-3)

      - Justification of conformity to new CS or alternative solutions per MDR Annex I

    - If required, a CAPA shall be opened to ensure full and timely implementation of actions required to maintain compliance.

    - All findings and actions shall be reported as part of Management Review (QSP 5.6-1) under the "Regulatory Updates and External Influences" agenda item. This ensures executive oversight of any regulatory-driven changes to the risk profile or technical documentation.

    - RA shall also notify the Document Control Administrator for updates to controlled documents such as QSP 9.1-3 (Technical Documentation – MDR), GSPR Checklists, and other quality system procedures that may be impacted.

    - Training shall be assigned where applicable to ensure affected personnel understand any changes to regulatory requirements, device risk, or associated procedures.

- PROCESS FLOWCHART

# Appendix 1 – Example Ranking System for Failure Mode and Effects Analysis

# SEV – Effects of failure (Severity)

| Ranking | Definition | Use and Design Clinical Effects; Process End Effects |
| --- | --- | --- |
| 5 | Catastrophic | Serious (irreversible) injury or death of the patient or user; very severe negative effect on the environment. |
| 4 | Critical | Serious (reversible) injury to the patient or user; severe negative effect on the environment. For example, this includes the need for a more invasive procedure such as surgery or increases case complication to fully treat the injury. Note: Any labeling issues that could lead to a recall must be ranked at a minimum of 4. |
| 3 | Moderate | Moderate injury to the patient or user; moderate negative effect on the environment. Decline of product performance or user confidence in the product/ company (e.g. customer is very annoyed or dissatisfied). For example, this includes actions taken to treat the patient or user within the scope and type of treatment already in progress. |
| 2 | Minor | Minor injury to the patient or user; minor negative effect on the environment. Slight decline of product performance or user confidence in the product/ company (e.g. customer slightly annoyed or inconvenienced). For example, this includes prolonging or delays in clinical procedure that do not pose a risk of greater injury to the patient. |
| 1 | Negligible / Cosmetic | No / virtually no injury to the patient or user; no / virtually no negative effect on the environment. No impact on product performance or user confidence in the product/ company; user may or may not even notice the failure. |

# OCC – Probability of Occurrence

| Ranking | Definition | Frequency |
| --- | --- | --- |
| 5 | Up to 1 in 20 (5 %) | Very High |
| 4 | < 1 in 20 to > 1 in 400 (0.25 to 5 %) | High |
| 3 | < 1 in 400 to > 1 in 15,000 | Moderate |
| 2 | < 1 in 15,000 to > 1 in 150,000 | Low |
| 1 | < 1 in 150,000 | Remote |

# DET – Probability of detection

| Ranking | Definition | Design |
| --- | --- | --- |
| 5 | Remote | Current controls are very unlikely or unable to detect potential cause and subsequent failure mode or no controls are identified. |
| 4 | Low | Low chance that current controls will detect the cause and subsequent failure mode. |
| 3 | Moderate | Moderate chance that current controls will detect the cause and subsequent failure mode. |
| 2 | High | High chance that current controls will detect the cause and subsequent failure mode. |
| 1 | Very High | Very high chance that current controls will detect the cause and subsequent failure mode. |

# Determination of acceptable risk:

# Risk Priority Number (RPN) = Effects of failure (severity) X Probability of Occurrence X Probability of detection

# Risk Index

| Risk Level | RPN Range | Acceptability |
| --- | --- | --- |
| 1 | 1 through 9 | Broadly acceptable, risk has been reduced as far as possible (AFAP). |
| 2 | 10 through 27 | Risk should be reduced as far as possible (AFAP) with technical practicability and are balanced against benefit-risk. Risk is considered acceptable. Economic considerations are not acceptable to stop risk reduction. |
| 3 | 28 through 59 | Risk is unacceptable and should be reduced as far as possible (AFAP); technical practicability is balanced against benefit-risk and risk is reduced even at considerable costs. |
| 4 | > 60 | Intolerable, risk is unacceptable and must be reduced. |

Appendix 2 – Sample Outline of Risk Management Plan

# Risk Management Plan

For

(Insert product name and model numbers here)

- Scope

  - Device

LIST PRODUCT AND INCLUDE DESCRIPTION FROM 510(k), Technical File, etc.

  - Life Cycle

SAMPLE – Life cycle of this device is from ‘development through obsolescence’.  Risk Analysis will be performed in accordance with QSP 7.3-1 Design and Development and QSP 7.3-3 Risk Management during the development process and transfer to manufacturing.  Post-production monitoring will be performed in accordance with QSP 7.3-3 Risk Management and QSP 9.1-7 Post Market Surveillance.  We will perform and document an annual audit of the risk management documents.

  - Risk Management documents

SAMPLE – The Risk Management documents include but are not limited to the following documents:

- Risk Management Plan

- Form 7.3-3-1 Risk Management Report and Summary [OR INSERT OTHER APPLICABLE DOCUMENT]

- Form 7.3-3-2 FMEA [OR INSERT OTHER APPLICABLE DOCUMENT]

- Risk Ranking Table

- Post-production monitoring documents or reference to those documents

- Revisions to:

      - Risk Management Plan,

      - Risk Management Report & Summary and

      - FMEA

      - Annual reviews, including any changes to the Risk Management Documents

- References (example – change as needed)

  - QSP 7.3-1		Design and Development

  - QSP 7.3-3 		Risk Management

  - Form 7.3-3-1		Risk Management Report and Summary

  - Form 7.3-3-2		FMEA

  - ISO 14971:2012	Medical devices – Application of risk management to medical devices

# 3.0	Responsibility

| Person | Responsibility |
| --- | --- |
|  | SAMPLE – Project Leader, Development, FMEA |
|  | SAMPLE – Risk Management Plan, FMEA, Change Control, Complaints, Internal Audits, Management Review |
|  | SAMPLE – Design, FMEA |
|  | SAMPLE – FMEA, Design Team Member, Customer Feedback, Post Market Surveillance |
|  | SAMPLE – FMEA, Design Team Member, Customer Feedback, Post Market Surveillance |

- Criteria for Acceptability of Risk

  - Risks will be evaluated in accordance with QSP 7.3-3 Risk Management for:

SAMPLE - Effects of Failure, Probability of Occurrence, and Probability of Detection

  - Acceptable Risk per RPN for [PRODUCT] is XX.

  - Any RPN’s above these values need to have actions taken to reduce the risk.

- Requirements for Review of Risk Management Activities

  - A Failure Mode and Effects Analysis Report will be initiated during product design and updated as necessary.

(List Requirements here) – SAMPLE:

- Device characteristics that could impact on safety (reference Annex C in ISO 14971) will be identified in the on Form 7.3-3-1 Risk Management Report and Summary or via other means.

- The FMEA report for the product analyzes product and use hazards.

- A FMEA will be completed for both hardware and software.  Data on the FMEA Table includes:

  - Process/Item Function Process/Function Description

  - Potential Failure Modes (Identification of Hazards)

  - Potential Effect(s) of Failure & Estimation of Risk

  - Potential Causes(s) or Mechanism(s) of Failure and Estimation of Risk

  - Current Controls and Estimation of Detection

  - RPN value (Risk Evaluation)

  - Recommended Action(s) if RPN exceeds acceptable limits – including Responsibility and Target Completion Date; Action Taken; Severity, Occurrence & Detection values and revised RPN.

# Note:	 If RPN still exceeds acceptable limits, further action will be required.

5.2	The will identify whether or not a new risk has been created.  If so, this risk will be added to the FMEA table.

5.3	The Risk Management Report and Summary [or list other applicable documents] will identify / summarize whether or not the RPN numbers meet the values in the Risk Management Plan and will include a statement that  all known hazards will be identified.

- Verification Activities

  - Verification and validation testing according to QSP 7.3-1 Design and Development will be cross referenced in the Current Controls and/or the Recommended Actions columns of the FMEA table.

7.0	Approvals

7.1	The Risk Management Plan must be reviewed and approved prior to start of the FMEA (risk analysis).

7.2	The Risk Management Report and Summary (Form 7.3-3-1) or other applicable document (list here) will be reviewed and approved to ensure completeness and conformance to QSP 7.3-3 Risk Management procedure and this Risk Management Plan.

Note:	Approvals are required by the President and Management Representative.

# 8.0	Location of Risk Management Documents

  - The Risk Management documents will be located in the Technical File binder for the product family.