Tulip Medical

Quality System Procedure

QSP 7.3-5

Ethylene Oxide (EtO) Sterilization Process Control

Approvals

V.P Regulatory Affairs     							Date

Regulatory Affairs Coordinator 				                                    Date

Document Control Administrator					                  Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 | 10/24/2019 | Initial release. |
| 1 | 5/28/2025 | Added references to Regulation (EU) 2017/745 (MDR) and relevant MDCG and EN ISO guidance documents to support EU regulatory compliance. Inserted clarifying note in the Scope section to reflect alignment with MDR General Safety and Performance Requirements. |

- PURPOSE

To outline the procedure for sterilization processing and product release to ensure a product Sterility Assurance Level of 10-6 or less.

- SCOPE

This procedure applies to all terminally sterilized products manufactured by Tulip Medical Products that have a validated EtO sterilization process. All sterilization validations will be performed following appropriate industry guidelines (ISO/EN/AAMI).

Comprehension of, and satisfactory expertise in, the material covered by this Quality System Procedure by each employee will be indicated through the completion of all applicable and appropriate training requirements, as is indicated by the completion of Form 6.2-1-2 Training Record in accordance with QSP 6.2-1 Training and Qualification.

This procedure supports compliance with the applicable requirements of the EU Medical Device Regulation (EU) 2017/745

- REFERENCE DOCUMENTS

| Regulation (EU) 2017/745 – Medical Device Regulation (MDR), Annex I, General Safety and Performance Requirements MDCG 2020-16 Rev.1 – Guidance on sterilization validation (where applicable) EN ISO 14937 – Sterilization of health care products — General requirements for characterization of a sterilizing agent and the development, validation and routine control of a sterilization process for medical devices ANSI/AAMI/ISO11135:2014 – Sterilization of health-care products. Ethylene oxide. Requirements for the development, validation and routine control of a sterilization process for medical devices. |  |
| --- | --- |
| AAMI TIR28:2009 – Product adoptions and process equivalence for ethylene oxide sterilization. |  |
| EN ISO 10993-7:2008 – Biological evaluation of medical devices. Ethylene oxide sterilization residuals. |  |
| ISO10993-1:2009 – Biological Evaluation of Medical Devices. Part 1: Evaluation and testing within a risk management process. |  |
| Tulip Medical Product Quality Manual QM |  |
| Form 6.2-1-2 | Training Record |
| Form 7.3-5-1 | Lot Release Checklist |
| QSP 4.2-1 | Document Control |
| QSP 4.2-2 | Control of Quality Records |
| QSP 6.2-1 | Training and Qualification |
| QSP 7.3-1 | Design and Development |
| QSP 7.4-1 | Purchasing and Supplier Evaluation |
| QSP 7.6-1 | Control of Measurement Equipment. |
| QSP 8.3-1 | Control of Non-Conforming Product |
| WI 1.1 | Incoming Product Inspection Disposable Cannula Components and Accessories |

- DEFINITIONS

Bioburden – Population of viable microorganisms on or in the product and/or sterile barrier system.

Document Control – Consists of Regulatory Affairs (RA) and Quality Assurance (QA).

EO Residuals – Ethylene Oxide residuals left over on the device post sterilization.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (QSPs).

Pyrogen (LAL) – A fever producing substance.

Sterile – Free from viable microorganisms.

Sterility Assurance Level (SAL) – Probability of a single viable microorganism occurring on an item after sterilization.

Sterilization – Validated proves used to ender product free from viable microorganisms.

Sterilization Validation – The establishment of documented evidence that provides a high degree of assurance that a specific sterilization process will consistently result in a product that meets a pre-determined SAL.

- ABBREVIATIONS

DC – Document Control

MR – Management Representative

QA – Quality Assurance

QSP – Quality System Procedure

RA – Regulatory Affairs

- RESPONSIBILITIES

The Management Representative (MR) is responsible for maintaining this procedure in a current state.

Quality Assurance (QA) is responsible for overseeing the activities described in this procedure.

Quality Assurance (QA) is responsible for preparing, assembling and reviewing process documentation associated with this procedure.

Operations and/or Quality Assurance (QA) are responsible for the assembly of routine sterilization loads for shipment.

Note: In the case that products are packaged and labeled by Contract Manufacturer (Eagle Labs), the preparation of sterilization load will be carried out by Eagle Labs.

- PROCEDURE

  - Product preparation

    - Prior to sterilization, product will be assembled according to appropriate manufacturing procedures.

  - Sterilization

    - Non-sterile product will be transported to the approved contract sterilizer following any special shipping instructions.

    - Product shall be scheduled for sterilization in a validated loading pattern and process cycle specification. The process cycle specification indicates:

      - Number of Biological indicators.

      - Amount of process exposure (preconditioning time, EtO exposure time, aeration time).

      - Ensure the proper amount of EtO residuals for proper validation of EtO.

      - Sterilization process.

      - Loading pattern / configuration.

    - The contract sterilizer will handle and process the product according to an established validated sterilization process per supplier agreement.

    - The contract sterilizer and / or contract manufacturer will return the sterilization load to Tulip Medical Products along with the process results.

  - Sterilization Lot Release

    - Once the load is returned, the product will be placed in a quarantine area until all required processing information is reviewed and testing information is received.

    - For each completed product lot, the Lot History Record will be reviewed by QA. This review will include verifying acceptable product packaging, product quantities and completeness of Lot History Record documentation. Depending on the specific product and associated sterilization process, the following will be verified:

      - For EtO sterilized products, the parameters must be within the minimum and maximum validated range per the cycle specification card. Final release will be pending sterility results of the biological indicators in which all indicators in the load must have negative sterility (no growth).

      - All pyrogen test samples must demonstrate non-pyrogenicity.

      - The results of this inspection will be documented on Form 7.3-5-1 Lot Release Checklist.

    - The sterilization in process parameters for each batch will be recorded.

    - All testing results will be included in the respective Device History Record for the products sterilized.

    - Once the sterilization load is released, it will be moved from the quarantine area to the finished goods area.

    - All non-conformances related to sterilization processing or associated testing will be handled using QSP 8.3-1 Control of Non-Conforming Product. The sterilized product will remain in quarantine until all process or testing issues are resolved.

  - Product Re-Sterilization

    - Product re-sterilization is allowed if the number of times product can be re-sterilized has been validated to ensure there is no deterioration to the product and package performance.

  - Maintenance of Validation

    - The Sterilization Validation will be maintained per industry guidelines, using harmonized standards (EN) when appropriate. This will include product bioburden testing and environmental monitoring of manufacturing.

    - For commercially distributed products, sterilization process will be re-evaluated every 24 months following requirements defined in the appropriate standard.

    - Incorporate all design changes during re-evaluation of the EtO Sterilization process.

  - Quality Assurance will review new products and product changes as they are introduced and determine whether they can be included into the existing validated process cycle. This review will consider the following:

    - Product and package design.

    - Component sources (suppliers).

    - Materials.

    - Manufacturing processes.

    - Manufacturing environments.

- PROCESS FLOWCHART

  - N/A