Tulip Medical

Quality System Procedure

QSP 7.5-3: Product Labeling

Approval

V.P. Regulatory Affairs 	Date

Regulatory Affairs Coordinator						Date

Document Control Administrator 	Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 0 |  | Released documentation – Implementation of document for compliance to FDA CFR Part 801, FDA QSR Subpart K 820.120, CMDR, Medical Device Directive 93/42/EC and ISO 13485:2003 7.5.1.1 |
| 1 |  | Updated references to ISO 13485:2012 |
| 2 |  | Removed Appendix A and integrated it into the system as RFD-033. Section containing label samples removed due to the label examples being inaccurate. Section on in-house labeling operations updated to reference applicable work instructions. |
| 3 | 10/01/2019 | Added the ‘Date’ column to the Rev changes on this page. |
| 4 | 02/17/2020 | Correction on Section 7.9 In House Labeling Operations. 7.9.2 Work Instruction 3.2 to 2.2 Packaging, labeling and Final Inspection of Reusable Cannulas. Section 7.9.3 correction Work Instructions 5.2 to 3.2 Packaging, Labeling and Final Inspection – Sterile Single- Use |
| 5 | 2/23/2021 | Add to reference – WI 3.2-1 and WI 3.2-2. Add WI 3.2-1, and WI 3.2-2 to section 7.9.3 |
| 6 | 03/26/2025 | 3.0 Updated Reference documents 7.6 Updated steps in label inspection 7.7 Added step to when updating an IFU 7.9 Added Change request when updating 7.10 Added Document control Removed In House Labeling process 8.0 Added Flowchart |
| 7 | 5/23/2025 | Updated all applicable references from MDD (93/42/EEC) to MDR (EU 2017/745) throughout the procedure body. Ensured alignment with Article 27 and Annex VI MDR for UDI system, Article 10(11) for translations Clarified labeling requirements per MDR Annex I and Chapter III. |

- PURPOSE

To ensure product labeling contains the required information for compliance with FDA medical device regulations, ISO 13485, Health Canada and the Regulation (EU) 2017/745 on medical devices (MDR).

- SCOPE

This procedure applies to all devices manufactured in-house or by contract manufacturers.

This procedure is to be followed by all Tulip Medical Products personnel. Comprehension of, and satisfactory expertise in, the material covered by this Quality System Procedure by each employee will be indicated through the completion of all applicable and appropriate training requirements, as is indicated by the completion of Form 6.2-1-2 Training Record in accordance with QSP 6.2-1 Training and Qualification.

- REFERENCE DOCUMENTS

| Food and Drug Administration 21 CFR 820 Quality System Regulation |  |
| --- | --- |
| International Standard Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes ISO 13485:2016 |  |
| International Standard Medical Devices – Symbols to be used with medical device labels, labeling and information to be supplied – Part 1: General requirements ISO 13485:2016 |  |
| EN 1041:2008 – Information supplied by the manufacturer of medical devices |  |
| Regulation (EU) 2017/745 on medical devices (MDR). |  |
| Canada’s Medical Device Regulations SOR/98-282 |  |
| Form 4.2-1-2 | Change Request |
| Form 4.2-1-4 | Document Review Checklist for IFU |
| Form 6.2-1-2 | Training Record |
| Form 7.5-3-1 | Label Approval |
| QSP 4.2-1 | Document Control |
| QSP 4.2-3 | Change Control |
| QSP 6.2-1 | Training and Qualification |
| QSP 7.3-1 | Design and Development |
| QSP 7.3-2 | Device Master Record |
| QSP 7.4-1 | Purchasing and Supplier Evaluation |
| QSP 7.4-2 | Product Inspection |
| QSP 9.1-4 | Canadian Medical Device Regulations |
| RFD-033 | Language Requirements for Medical Devices in Europe |
| RFD-106 | Class II References |
| RFD-112 | Class I References |
| RFD-114 | Class II Approved Labels |
| RFD-118 | Class I Approved Labels |
| WI 3.2-2 | Creating & Managing Single-Use Device GTIN |

- DEFINITIONS

Labeling – All labels and other written, printed or graphic matter accompanying or attached to the device or its container. The term accompanying extends to tags, pamphlets, circulars, booklets, Instructions for Use, etc. that may be included with the device.

Management Representative – The member of Tulip Medical whose cross-functional role encompasses the development and monitoring of the quality management system to ensure accuracy and effectiveness of all Quality System Procedures (QSPs).

UDI – Unique Device Identifier

- ABBREVIATIONS

CMDR – Canadian Medical Device Regulations

FDA – Food and Drug Administration

IFU – Instructions for Use

QSP – Quality System Procedure

QSR – Quality System Regulation

- RESPONSIBILITIES

The Management Representative (MR) is responsible for maintaining this procedure in a current state and for ensuring all product labeling complies with this procedure.

Production Personnel who inspect product are responsible for reviewing labeling as described in this procedure.

- PROCEDURE

  - Tulip Medical reviews and approves the content of labeling and ensures labeling complies with regulatory requirements. See QSP 4.2-1 Document Control, QSP 7.3-1 Design and Development and QSP 7.3-2 Device Master Record.

  - Labels are designed, printed and applied to:

    - Define the device

    - Remain legible during the conditions of processing, storage, handling, distribution and use

    - Prevent mix-ups during handling and storage

  - Labels contain the required information identified in the labeling requirements of the MDR, FDA QSR and other FDA labeling regulations, CMDR and ISO 13485.

  - Sources for labeling requirements include, but are not limited to:

    - FDA: Refer to 21 CFR Part 801 for required information.

Note:	FDA labeling guidance is available on FDA’s website, and some devices may have specific labeling requirements.

- Canada: Refer to CMDR Section 21, 22 and 23.

Note:	CMDR may have specific requirements for some devices on their website.

- MDR: Refer to chapter III and other applicable Articles and Annexes in the MDR.

- International and National Standards applicable to the product may also require specific labeling or symbols to be placed on the packaging, label, or the device itself.

Note: 	Other standards or guidance documents may also have relevant labeling requirements.

  - Where applicable, for products sold in the European Union, graphical symbols must be used in accordance with:

    - Where applicable, refer to ISO 15223 for additional symbols or information to be provided on the device and in the Instructions for Use (IFU).

    - Where applicable, refer to EN 1041 for additional information to be provided on the device and in the IFU.

  - Label Inspection

    - During Final Inspection (see QSP 7.4-2 Product Inspection), label inspections are conducted by Production Personnel to ensure correct label was applied to the product.

    - All Final Product labels must have Form 7.5-3-1 Label Approval, to ensure accuracy of information is up to date

    - Refer to RFD-118 for Class I Approved labels

    - Refer to RFD-114 for Class II Approved labels

    - A GTIN number will be required to be made when creating a new product’s label

    - See Work Instruction 3.2-2 when creating a GTIN number

    - RFD-112 references Class I approved GTIN

    - RFD-106 references Class II approved GTIN

    - UDI System Requirements per Article 27 and Annex VI MDR

- To be registered in EUDAMED, where applicable

  - The IFU should be developed in English and accepted by the Notified Body and/or Health Canada, as required.

    - A Document Review Checklist Form 4.2-1-4 must be completed to ensure clarity and completeness.

  - Translations

    - The labeling must be translated into the languages of the countries we are selling to, as stated in the labeling requirements of CMDR, the European Union and elsewhere, as applicable, unless otherwise justified by documented rationale. If translations are required, the labeling materials will be translated to the desired number of languages depending on the countries prior to placing product in the markets.

Note:	In Canada, the labeling may need to be provided in French and English. Refer to QSP 9.1-4 Canada Medical Device Regulations.

    - In accordance with MDR 2017/745 Article 10(11) and Annex I, most Member States within the European Union (EU) and European Economic Area (EEA) require that device labeling, including packaging and Instructions for Use (IFU), be provided in their official national language(s) when the product is placed on the market in those countries. Additional language requirements may be communicated by the customer and/or distributor. Tulip Medical addresses any related risks through a documented rationale and product-specific risk assessment to ensure safe and effective use by the intended user.

      - RFD-033 Language Requirements for Medical Devices in Europe can be used as an initial guide for the language requirements for medical devices in Member States of the European Union or participants in the European Economic Area Treaty.

Note:	Prior to translating, verify languages listed in RFD-033 Language Requirements for Medical Devices in Europe to ensure they are still current.

    - An approved supplier must perform translations of the labels; preferably with experience in medical device translations to ensure the meaning of the label information is not inadvertently modified. See QSP 7.4-1 Purchasing and Supplier Evaluation for information on supplier approval.

    - Translations are recommended to be either certified by a translating service, or verified by an independent source, such as another person within the translator service provider, distributor, etc.

Note:	In some cases, a distributor or other qualified source may perform translations. It is recommended that the translated documents be reviewed/verified by an approved translation service. Documentation of this review/verification will be maintained.

  - Any updates made to an IFU/Product label must be completed with a Form 4.2-1-2 Change Request per QSP 4.2-3

  - DCA to log and record all updates made for IFU/Product labels per QSP 4.2-1

- FLOWCHART