Tulip Medical

Quality System Procedure

QSP 8.2-2: Customer Feedback and Complaints

Approval

V.P. Regulatory Affairs 							Date

Regulatory Affairs Coordinator  						Date

Document Control Administrator 						Date

| Rev | Date | Description of Change |
| --- | --- | --- |
| 5 |  | Combine Form 8.2-2-2 Complaints and Form 8.2-3-1 RMA into a single form. |
| 6 |  | Incorporate Customer Survey cards. |
| 7 |  | Updated reference to QSP 9.1-5 and deleted reference to Netsuite (old inventory control software). |
| 8 |  | Complaint and RMA forms separated. References updated to include new forms. Procedure updated to match the current procedure used by Tulip Medical for processing complaints. Included provision in managing Complaints from OUS (Outside US) Countries in Sec. 7.7.4. |
| 9 |  | Title change from ‘Customer Complaints’ to ‘Customer Feedback/Complaints.’ Simplified and updated procedure in response to new Customer Feedback procedure and separation of RMA and Complaint procedures. Addition of section on servicing, as well as elaboration on warranty information, monthly meetings, and the general procedure. Included specification of the threshold for a Complaint, as well as information on how to determine if an investigation needs to be performed. |
| 10 |  | Updated the following sections: Sec. 1 & 2: added clarity Sec. 3: added required reference required doc and removed those that were not needed Sec. 6: Clarified responsibilities Sec. 7: Added Documentation Review section as part of “Complaint Evaluation and Investigation. Added “Reply to Complainant” section, as well as requirements for records. Added Attachment 1, 2, 3 flow charts for complaint determination in the US, EU, and Canada. |
| 11 | 10/10/2019 | Added the ‘Date’ column to the revision chart Updated the following sections: - Eliminate “forms” from the items supplied to the customer. - Added “Feedback” before complaints. - Added, “ Form 8.2-2-1 RMA form” for the documents required. - Added, "Feedback definition"; added 'non-Tulip products' to the last paragraph of Warranty - Added, “DCS – Document Control Administrator” - Changed responsibility of ‘Operations Manager’ to ‘Client Care Specialist’ - Paragraph 4, updated paragraph. 7.1.3 Open a CAPA – to clarify how and when a Complaint initiates a CAPA. 7.2.6 added ‘(see section Opening a CAPA)’ |
| 12 | 12/17/2019 | Correcting DCS – Document Control Specialist to DCA – Document Control Administrator |
| 13 | 05/18/2020 | Changed Title to Customer Feedback ‘and’ Complaints - Purpose: updated the purpose; moved Note to Procedure - Scope: updated what and whom of the scope - Reference Documents: removed Medical Device Regulations, referred to in QM and not necessary in QSP; removed Form 8.2-2-2-2; added Form 8.5-1-2; added Form 8.5-1-1; removed RMA Form; added WI 4.2; - Definition: changed Feedback; updated Management Representative; updated Warranty Period from Tulip website - Abbreviation: added QM – Quality Manual; added EVP; - Responsibilities: updated DCA from RFD-031; changed to ‘All Company Personnel’ and removed CCS; updated MR; updated QA and RA - Procedures: update the Note; added the last paragraph of origination; 7.1 added source paragraph; 7.1.1: added Form 8.2-2-7; restructured; changed Phone Calls received by any Company personnel; removed notes and added to each way of receiving; 7.1.2 restructured paragraph - Feedback is logged and coded for the nature of the feedback (FORM 8.2-2-8); 7.1.3 restructured paragraph - All feedback is analyzed to determine if a response action is necessary.; added 7.1.3.1 The EVP approves a return or replacement product (QSP 8.2-3; WI 4.2); |
| 14 | 6/29/2020 | 4.0 Definition. Corrected the “Complaint” definition to the FDA’s definition. |
| 15 | 03/20/2025 | 3.0 Reference Documents -Corrected the references on the following Forms: 8.2-2-2, 8.2-2-5, 8.2-2-6, 8.2-2-7 -Addition of the following Forms: 8.2-2-0, 8.2-2-1, 8.2-2-3 -Removed WI 4.2 4.0 Added External Party 5.0 EVP is updated to VP 6.0 Removed Sponsor 8.0 Updated and rephrased procedure 11.0 Updated Flowchart Attachment 1 Updated Attachment 2 Updated All sections - QA/RA Manager is the QA Manager -RA is Regulatory Affairs where applicable -Removed Dropbox from the entire document. -Reformatted subparagraphs -Cross-referenced to Customer questionnaire |
| 16 | 5/23/2025 | Updated EU vigilance reporting references: Replaced obsolete MEDDEV 2.12-1 Rev. 8 references in Attachment 2 with MDR (EU) 2017/745, Articles 87–88 and supporting guidance from MDCG 2023-3 and MDCG 2023-4. Clarified criteria for “serious incidents” per MDR Article 2(65). Revised Attachment 2: EEA and Switzerland Decision Tree: Rewritten to align with MDR terminology and decision criteria for reportability. Reference Documents Section (EU Subsection): Added: Regulation (EU) 2017/745: Articles 87, 88, MDCG 2023-3: Serious Incident Determination – Decision Tree, MDCG 2023-4: Manufacturer Incident Report Timelines and Responsibilities |
| 17 | 08/07/2025 | Updated QSP from MR to PRRC Section 3.0 Added MDR (EU) 2017/745 Updated correct section for Canada 57-59 Section 6.0 Updated responsibility for a PRRC |

- PURPOSE

The purpose of this procedure is to define Tulip Medical’s requirements in gathering information from customer feedback and complaint handling.

- SCOPE

This procedure applies to Tulip Medical products distributed in the United States, Canada, European Union, Brazil, Australia, and Japan, and is the primary document meeting the applicable regulatory requirements for gathering customer feedback and complaint handling as defined in Tulip’s Quality Manual (QM 1).

The Vice Presidents and Quality Assurance Manager are trained on the intake of complaints, and all other Company personnel are trained on the intake of customer feedback.

- REFERENCE DOCUMENTS

| MDCG 2023-3 | Serious Incident Determination – Decision Tree |
| --- | --- |
| MDCG 2023-4 | Manufacturer incident report timelines and responsibilities |
| MDR (EU) | 2017/745 |
| USA | FDA 21 CFR 820.198 |
| EU | ISO 13485:2016: 4.21, 7.2.3, 8.2.2; 8.2.3 |
| Australia | TGA 2002 Sch3 Part 1 Clause 1.4(3)c(i) |
| Brazil | RDC ANVISA 67/2099 – Art. 5, 6, 7, 8, 9 |
| Canada | SOR/98-282, Section 57–59 |
| Japan | PMD Act Article 228-20-2 MHLW MO169: 62.6 |
| Form 8.2-2-0 | Complaint Form – Customer |
| Form 8.2-2-1 | Complaint Form – Distributor |
| Form 8.2-2-2 | Complaint Form – Distributor’s Customer |
| Form 8.2-2-3 | Complaint Investigation |
| Form 8.2-2-4 | MDR Determination - FDA |
| Form 8.2-2-5 | MDR Determination – Canada and Australia |
| Form 8.2-2-6 | MDR Determination – Brazil, Japan |
| Form 8.2-2-7 | MDR Determination – EU |
| Form 8.2-2-8 | Customer Feedback |
| Form 8.5-1-1 | CAPA Log |
| Form 8.5-1-2 | Corrective and Preventive Action |
| QSP 5.6-1 | Management Review |
| QSP 8.2-3 | Returns & Exchanges |
| QSP 8.5-1 | Corrective and Preventive Action |
| QSP 9.1-2 | Conducting Recalls – US FDA |
| QSP 9.1-5 | Recalls, Corrections, and Removals |
| QSP 9.1-6 | Medical Device Reporting |
| Website | https://www.tulipmedical.com/contact-us/ Tulip Products Limited Warranty |

- DEFINITIONS

Complaint – Any written, electronic, or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety, effectiveness, or performance of a device after it is released for distribution.

Feedback – Customer satisfaction and opinion, comments, and expression of interest in a product, a service, or a complaint-handling process.

Use-Error – Act, or omission of an act that results in a different medical device response than intended by the manufacturer or expected by the user.

External Party Contributor – An outside entity that contributed to the complaint

- ABBREVIATIONS

PRRC – Person Responsible for Regulatory Compliance

QSP – Quality System Procedure

RA – Regulatory Affairs

VP – Vice Presidents

- RESPONSIBILITIES and AUTHORITIES

All Company Personnel are responsible for ensuring that any items returned under a complaint are properly identified and directed to QA for investigation.

Document Control Administrator (DCA) is responsible for the maintenance of company documents. He/she controls the numbering, sorting, filing, storing and retrieval of both electronic and hard copy documents produced by Regulatory, Management or by departments within Tulip Medical Products.

Vice President of Finance (VP) is authorized to approve Return Merchandise Authorizations (RMA) and replacement orders QSP 8.2-3

Person Responsible for Regulatory Compliance (PRRC) The PRRC ensures Tulip Medical’s complaint handling process complies with MDR Article 15, including regulatory determination, timely reporting under Articles 87–88, oversight of technical documentation, and interface with Competent Authorities when required. In addition, the PRRC ensures the preparation of customer feedback and complaint reporting and analysis for the bi-yearly Management Review meetings.

Quality Assurance (QA) is responsible for producing quality documentation, maintaining quality management systems, carrying out product quality reviews to ensure that all devices comply with regulatory requirements.

- GENERAL INFORMATION

  - All Company personnel is responsible for notifying QA if they become aware of product experiences that could potentially be classified as a product complaint or customer feedback.

  - All information sources foreign or domestic, are considered; including information derived from commercial marketing experience, post-market clinical investigations, post-market surveillance studies, reports in the scientific literature, and unpublished scientific papers.

  - QA ensures that Distributors are responsible in accordance with a Distribution Agreement to notify the Company of any reported Adverse Event or Product Complaint associated with the products.

  - Warranty Period – Tulip Medical Products warrants its reusable cannulas and accessories, and all parts thereof, to be free from defective materials and workmanship for the period of one year from the date of original purchase. This warranty covers parts and labor. (Website)

  - Product experience information is to be communicated to QA in a written Complaint Questionnaire Form based on the customer. One questionnaire will be assigned to each item.

Note:	Tulip Medical Products are very durable surgical instruments and, under conditions of normal use, generally remain undamaged well beyond the warranty period, refer to Tulip’s website under Support. However, as with any surgical instrument, potential misuse or mishaps can occur. Damaged surgical instruments should never be used on a patient. Should any Tulip Medical Product become damaged, always recommend

- The customer stop uses immediately

- If possible, return it to Tulip Medical for evaluation

- Instruments that are exposed to surgery or used on a patient must be cleaned and decontaminated by the customer prior to returning the product for evaluation

- Before a customer returns one or more of our products, supply the customer with proper authorization and handling information. (QSP 8.2-3)

  - Customer Feedback and Complaints can originate from many sources and be submitted via email, online contact forms, telephone, social media, and online product reviews.  Often the complaints will be submitted directly from customers to any Company personnel and recorded on a Complaint Questionnaire Form

  - VP reviews all reported product experiences and determines whether it constitutes a complaint.

  - Complaints and Customer Feedback is reviewed at the QA Weekly Feedback Meeting.  Minutes are recorded by the DCA and added to the complaint if applicable.

  - Investigations are required for any complaint pertaining to a concern with product quality, whether the item is a Tulip or Non-Tulip Item.

  - Outside of this requirement, investigations are left to the discretion of the QA, PRRC, and Operations.

  - Investigations are required for any complaint pertaining to a concern with product quality, whether the item is a Tulip or Non-Tulip item.

  - Additionally, complaints involving contributions from external parties must incorporate information from the Customer Questionnaire as appropriate, to identify any external factors.

  - All external communication related to the complaint must be documented and attached to the complaint packet for traceability.

  - The PRRC and Operations are responsible for ensuring these requirements are followed before closing the complaint investigation.

  - PROCEDURE

  - Customer Feedback Types:

- Product information inquiries

- Customer Service

- Cosmetic defects

- Delivery errors (delayed deliveries)

Note:	Delivery errors such as damaged packages and/or damaged products is a complaint as the Company tracks these vendor issues.

  - Customer Feedback - Processing

    - Feedback are recorded on Form 8.2-2-8.

    - Feedback is logged and recorded for the nature of the feedback and analyzed for reporting to Management Reviews.  QA will analyze trends provide the PRRC with a report at the conclusion of each calendar month. The PRRC will summarize the data and trends for presentation during Management Review Meetings QSP 5.6-1

    - Feedback is reviewed by RA to determine if a response action is necessary.

    - The VP approves a return or replacement product per QSP 8.2-3

    - Upon determination of response and completion of any associated tasks or updates, the feedback is considered ready to close through the signature of a member of Management. Form is scanned and filed.

- Complaint Types

There are two types of complaints: Incident-driven and Review-driven.

  - A Review of Data issues may be noticed as a trend in repeated failures of medical devices. This is a reliability problem and falls within the definition of a complaint.

  - An Incident-driven issues require immediate action. If the failure could have led to serious harm or death to a patient, this is an incident-driven issue.

  - Complaint Investigation Processing Form 8.2-2-3

Note:	Tulip will make three attempts in contacting the customer to complete their Complaint Questionnaire Form. If the customer has not responded after 30 days and three tries, Tulip will then close the complaint, sighting “no customer response.”

    - A completed questionnaire must be provided prior to commencement of investigation.

•	Form 8.2-2-0	Customer (Domestic)

•	Form 8.2-2-1	Distributor

•	Form 8.2-2-2	Distributor’s Customer

    - When a decision is made not to investigate, justification for this decision must be documented on the Complaint Investigation in the “Justification for No Investigation” section.

  - Part I Complaint Investigation - MDR Determination

    - Determine if the complaint is reportable, if not proceed to Part II

    - If the complaint is a ‘reportable’ event, select the appropriate form upon determination

- Form 8.2-2-4 MDR Determination - FDA

- Form 8.2-2-5 MDR Determination – Canada and Australia

- Form 8.2-2-6 MDR Determination – Brazil, Japan

- Form 8.2-2-7 MDR Determination – EU

    - For reporting guidance and procedures pertaining to Adverse Events refer to QSP 9.1-5

    - Time Reporting Requirement Table

| Requirement | Timeframe | Regulation/Source |
| --- | --- | --- |
| MDR Determination | Within 48 hours | CMDR, Form 8.2-2-3 |
| Canadian Complaint Notification | 10/30 calendar days | CMDR Section 58/8.4.3 |
| Complaint Investigation Initiation | Within 48 hours | QSP 8.2-2 |
| Complaint Investigation Completion | Within 30 calendar days | QSP 8.2-2 |
| Follow-up on Unresponsive Customer | 30 days (3 attempts) | QSP 8.2-2 Section 8.4 |

  - Complaint Investigation – Regulatory Investigation to be completed by RA or PRRC.

    - Part II Item Inspection

- Fill the required part information associated with the complaint

- Check the box that best pertains what was the evaluation based on; photo, returned product or N/A

- Always request photos emailed immediately or return of the complaint item, If available.

Note:	For the safety of Tulip personnel: If the complaint item was used and not returned in a biohazard bag as instructed, regardless of if it was sterilized, the complaint item will be safely and immediately discarded, and no investigation will commence.  The customer will be immediately notified of this action taken.

- Evaluate the complaint item by answering the bulleted questions.

- If tests were performed on the complaint item, include the test documents in the complaint packet.

    - Part III Risk Evaluation

- In the QA Weekly Feedback Meeting, the complaint item is evaluated based on the most corresponding Risk Table by the QA, RA, PRRC, VPO, and R&D. If discussed in the meeting, include the minutes in the complaint packet.

- Document Evaluation review number and revision

    - Part IV Corrective and Preventive Action (CAPA)

- Determine if the complaint needs to be elevated to a CAPA.

- This decision will be documented on Form 8.5-1-2 in accordance with QSP 8.5-1 Corrective & Preventative Action.

- The CAPA process is an approach to eliminating the cause of an existing or potential nonconformity, defect, or other undesirable situation to prevent the occurrence. The actions taken to eliminate the root causes of actual nonconformance(s) shall be to a degree appropriate to the magnitude of problems and consistent with the associated risks.

      - Implement a CAPA when there is documented evidence that:

- All proposed CAPAs are completed and reviewed with the appropriate affected function

- All CAPA actions have been verified/validated

- All relevant information regarding the investigation and root cause analysis is completed and reviewed with appropriate affected functions

Note:	Always check if a CAPA was previously addressed/opened with the same issue.

      - Closing a CAPA, refer to QSP 8.5-1.

    - Part V Recall

      - If a recall deems necessary, refer to the following

    - QSP 9.1-2 Conducting Recalls – US FDA

    - QSP 9.1-5 Recalls, Corrections and Removals

    - QSP 9.1-6 Medical Device Reporting

    - Part VI Device History Record

    - Fill in the required information for the item’s origin of build and receiving information

    - Determine if item is still under Tulips Warranty

    - Select and check box to indicate item type and issue

    - History documents are pulled and copied by the DCA and included with the complaint packet.

    - Part VII-A Complaint Investigation Summary

    - Attach appropriate completed Questionnaire

    - Select box if External Party Contributions were involved, if yes attach all documents

    - Summarize the complaint with as much detail of the investigation and add

Note:	If an investigation could not be performed within the allotted timeframe be sure to note the reason in the summary, refer to 10.0 Complaint Timeline in this QSP.

    - Part VII-B Complaint Investigation Disposition and Justification

    - Include Disposition (What) and Justification (Why)

    - Be as thorough in your summarization especially if Tulip’s decision is to not replace the complaint product.

    - Part VIII RMA or Replacement S.O.

    - There are two parts in this section:

      - Shipping Issues – Tulip tracks the vendor for shipping trends, i.e. damaged package, damaged product, incorrect delivery to the customer, etc.  If the complaint was only regarding the shipment, describe in detail and skip all other sections.

      - Replacement QSP 8.2-3 – Only the VP authorizes replacements and RMAs and places a new sales order.  Include a copy of the RMA and/or sales order in the complaint packet.  The VP will notify the customer of the RMA and shipping tracking number.

    - The Production/Shipping Associate will create a shipping label for the VP

    - A copy of the shipping label is included in the complaint packet

    - The sales order, ship label, and RMA (if applicable) is scanned and emailed to the customer. The pdf is filed in the associated folder.

    - Part IX Reply to Complainant

VP prepares an email reply to the complainants based on the investigation disposition and justification.  A reply to the complainant is NO, based on

    - The complaint item was not returned for evaluation and therefore, no findings, no evidence to suggest user error or abnormal use.

    - The complaint is a ship/delivery issue.

  - Complaint Investigation Check List

The DCA is responsible to review the complaint packet using the complaint checklist as a guide.

    - The DCA checks for signatures, dates, and inclusions then submits packet to the PRRC.

  - Final Review by PRRC

- Investigation summary

- Edits if necessary, if an edit is made the PRRC returns to QA Manager to correct the edits. QA Manager returns packet to DCA to proofread the edits and submits to PRRC.

    - The PRRC returns packet upon completion to the DCA who then checks the packet one last time

- Check for signatures and dates

- Scan entire complaint packet placing the pdf in the associated complaint soft-copy folder and files hard copies in the appropriate file cabinet.

- DCA will log the information into the Complaint Log. The log is then used to track the trend analysis.

    - Centrifuge Complaint

      - In the event that Tulip Medical Products is notified of a defective centrifuge within an active warranty period, initiation of a Complaint Questionnaire Form will be followed by Tulip Medical contacting the appropriate supplier to coordinate a request for servicing for the customer.

      - The QA Manager will request that copies of any documents or reports associated with the supplier’s investigation and/or servicing of the product be provided to Tulip Medical for inclusion with its Complaint Questionnaire Form

- Complaint Timeline

Fig. 1

- A few exceptions to number 1. Intake Personnel (48 hours)

- If more information is needed to start an investigation, the 48 hours would not comply. Make a minimum of three call back attempts to contact customer.

- If the question regarding a patient involvement was not asked when the complaint was received, then 3b. MDR Determination 24 hours would not comply until that information is given.

- Waiting for a complaint item to be returned or photos emailed to start the investigation; however, all history documentation can be collected prior to receipt, if the information is given.

- OVERALL PROCESS FLOWCHART

# Attachment 1: Product Complaint and US Medical Device Reporting Decision Tree

Product Feedback

No

Does the reported information reasonably suggest that a Company device may have caused or contributed to the DSI?

Does the reported information suggest that a patient who was treated with a Company device experienced a Death or Serious Injury (DSI)?

No

Justify for “No investigation needed”

Does the reported information meet the definition of a Product Complaint for a device distributed by the Company?

No

Yes

Yes

PRODUCT COMPLAINT

(requires US MDR if marketed in the US)

Yes

Does the reported information suggest that a Company device malfunctioned AND that the device or a similar device distributed by the Company would be likely to cause or contribute to a DSI if the malfunction were to recur?

Yes

No

PRODUCT COMPLAINT

(US MDR not applicable)

# Product Complaint Definition: Any written, electronic, or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety, effectiveness, or performance of a device after it is released for distribution (per 21 CFR 820).

# Attachment 2: EEA and Switzerland Report Decision Tree

# Per Regulation (EU) 2017/745, Article 87 and Article 88, and further guidance provided in MDCG 2023-3 and MDCG 2023-4, incidents meeting the definition of serious incidents must be reported to the relevant Competent Authority.

# Any Event That Is An INCIDENT, Must Be Reported to The Relevant Competent Authority

Did an EVENT occur and is the EVENT considered an INCIDENT per MDCG 2023-3

4.10?

No

Yes

Is the MANUFACTURER’s device suspected to be a contributory cause of the INCIDENT?

Not Reportable

No

Yes

Did the EVENT lead, or could have led to the death of a patient, user or another person?

OR

Did the EVENT lead to, or could have led to a serious deterioration in the state of health of a patient, user, or another person?

No

Yes

The Event Must Be Reported to The Relevant Competent Authority

# Attachment 3: Canada Mandatory Reporting Decision Tree